## Input
Device Type | Lot # | Identification # | Actuation Number | Rated Doses | Who Actuated? | Actuation Date | Method Name | Ellipticity | Inclusion | Area [mm^2] | Area % | COG Angle Diff [deg] | COG Mag Diff | COG x | COG y | 1Kg Mass PN | 1Kg Mass SN | 3Kg Mass PN | 3Kg Mass SN | AS Accel [mm/s/s] | Actuation Vel [mm/s] | Actuator PN | Actuator SN | Actuator Temp | Actuator Type | Approval Date - SP | Approval Date - SP | Approval Operator - SP | Approval Operator - SP | COM x | COM y | Dmax [mm] | Dmin [mm] | Ovality | CPU Speed | Cal. Target PN | Cal. Target SN | Camera Distance | Camera Height | Camera PN | Camera Power PN | Camera Power SN | Camera SN | Characterization Actuation Number | Characterization Mode | Characterization Time | Computer Mfg. | Computer Model | Contact Force [kg] | Custom ID #1 | Custom ID #2 | Custom ID #3 | Custom ID #4 | DAQ PN | DAQ SN | Device Creation Date | Drug Name | Encoder PN | End Time [ms] | End of Stroke Force [kg] | Final Delay [ms] | Frame Rate [Hz] | Frames | Hardware Version | Hold Time [ms] | I/O Cable PN | IAD [ms] | IMAQ Cable PN | IMAQ PN | IMAQ SN | Inclination [deg] | Initial Delay [ms] | Install Date | Iris Setting | LC Electronics PN | LC Electronics SN | Laser PN | Laser Position | Laser SN | Lens PN | Lens SN | Loadcell PN | Loadcell SN | Location | Major Length [mm] | Manufacturer | Measurement Date - SP | Method Creation Date | Method Creator | Method Description | Method Target | Method Version | Minor Length [mm] | Motor Driver PN | Motor Driver SN | Motor PN | Motor SN | Operating System | Pattern Distance [mm] | Perimeter [mm] | Power Supply PN | Power Supply SN | Profile Name | RAM | Return Accel [mm/s/s] | Return Vel [mm/s] | SP Operator | Service Tag | Shake Angle [deg] | Shake Delay [s] | Shake Diameter [mm] | Shake Distance [mm] | Shake Frequency [Hz] | Shake Motor Driver PN | Shake Motor Driver SN | Shake Repetitions | Shaking Motor PN | Shaking Motor SN | Software Version | Start Time [ms] | Stroke Length [mm] | System Part # | System Serial # | System Type | Test Comment | Threshold | Time Avg. Mode | Trigger Cable PN | Ver. Target PN | Ver. Target SN | Who Characterized?
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 13:11 | GW012-Mannul-SP30-MD-01 | 1.197 | 0.033 | 391.6 | 5.3 | 233.2 | 0.124 | -12.5 | -4.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -12.5 | -4.2 | 24.82 | 20.18 | 1.23 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 484 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 175.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.12 | TEVA | 2022-05-04 13:11 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.15 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.7 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 110 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 13:08 | GW012-Mannul-SP30-MD-01 | 1.124 | 0.045 | 425.4 | 5.7 | 220.8 | 0.259 | -12.6 | -3.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -12.8 | -3.1 | 25.22 | 21.62 | 1.167 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 310 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 123.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.35 | TEVA | 2022-05-04 13:09 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.66 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.82 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 58 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 13:06 | GW012-Mannul-SP30-MD-01 | 1.141 | 0.06 | 451.4 | 6.1 | 184.8 | 0.063 | -12.3 | -2.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.4 | MDx |  |  |  |  | -12.4 | -2.3 | 26.19 | 21.65 | 1.21 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 404 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 41 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.31 | TEVA | 2022-05-04 13:06 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.18 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.38 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 13:03 | GW012-Mannul-SP30-MD-01 | 1.084 | 0.044 | 432.8 | 5.8 | 193.4 | 0.068 | -11.4 | -2.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.4 | MDx |  |  |  |  | -11.5 | -2.3 | 25.06 | 21.7 | 1.155 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 362 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 117.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.13 | TEVA | 2022-05-04 13:03 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.27 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.56 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 54 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 13:00 | GW012-Mannul-SP30-MD-01 | 1.052 | 0.067 | 435.5 | 5.9 | 319 | 0.037 | -12.2 | -1.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.4 | MDx |  |  |  |  | -12.1 | -1.4 | 24.6 | 22.51 | 1.093 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 410 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 6.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.89 | TEVA | 2022-05-04 13:00 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.7 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.86 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:56 | GW012-Mannul-SP30-MD-01 | [8 empty] | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx | [9 empty] | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 268 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 |  | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 |  | TEVA | 2022-05-04 12:57 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 |  | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 |  | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 2 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:53 | GW012-Mannul-SP30-MD-01 | 1.127 | 0.051 | 352.2 | 4.8 | 328.2 | 0.092 | -13.2 | -3.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -13.1 | -3.6 | 23.38 | 19.5 | 1.199 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 394 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 139.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.16 | TEVA | 2022-05-04 12:53 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 19.67 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 67.36 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:50 | GW012-Mannul-SP30-MD-01 | 1.164 | 0.047 | 366.4 | 4.9 | 185.2 | 0.089 | -14.5 | -4.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -14.6 | -4.7 | 23.82 | 19.08 | 1.248 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 454 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 150.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.96 | TEVA | 2022-05-04 12:50 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 19.73 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.1 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 70 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:47 | GW012-Mannul-SP30-MD-01 | 1.127 | 0.047 | 456.1 | 6.2 | 228.9 | 0.147 | -11.8 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -11.9 | -1 | 26.49 | 21.87 | 1.211 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 402 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 114.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.3 | TEVA | 2022-05-04 12:47 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.45 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.17 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:44 | GW012-Mannul-SP30-MD-01 | 1.122 | 0.051 | 468.5 | 6.3 | 209.4 | 0.37 | -15.7 | -3.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -16 | -3.4 | 26.1 | 22.53 | 1.158 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 296 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 138.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.6 | TEVA | 2022-05-04 12:44 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.8 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.16 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:41 | GW012-Mannul-SP30-MD-01 | 1.059 | 0.052 | 448.7 | 6.1 | 179 | 0.231 | -11.5 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -11.8 | -1.2 | 25.63 | 22.41 | 1.144 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 408 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 75.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.32 | TEVA | 2022-05-04 12:41 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.96 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.86 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 82 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:38 | GW012-Mannul-SP30-MD-01 | 1.09 | 0.062 | 370.1 | 5 | 284.9 | 0.071 | -13.8 | -3.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -13.8 | -3.5 | 23.29 | 20 | 1.164 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 452 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 1.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.39 | TEVA | 2022-05-04 12:38 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.54 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.62 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:36 | GW012-Mannul-SP30-MD-01 | 1.154 | 0.038 | 451.3 | 6.1 | 67.9 | 0.043 | -12.1 | -2.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -12.1 | -2.5 | 26.29 | 21.78 | 1.207 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 372 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 172.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.43 | TEVA | 2022-05-04 12:36 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.46 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:33 | GW012-Mannul-SP30-MD-01 | 1.047 | 0.059 | 367.5 | 5 | 324.9 | 0.147 | -12.6 | -4.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -12.5 | -4.4 | 22.69 | 20 | 1.134 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 498 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 26.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 21.87 | TEVA | 2022-05-04 12:33 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.88 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.01 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 122 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 20 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:30 | GW012-Mannul-SP30-MD-01 | 1.159 | 0.048 | 502.3 | 6.8 | 210 | 0.348 | -13.4 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -13.7 | -1.1 | 27.83 | 22.74 | 1.224 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 418 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 153.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.9 | TEVA | 2022-05-04 12:30 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.2 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.57 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 100 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:28 | GW012-Mannul-SP30-MD-01 | 1.101 | 0.063 | 442 | 6 | 241.5 | 0.212 | -11.3 | -2.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -11.4 | -2.2 | 25.44 | 21.63 | 1.176 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 368 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 13 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.6 | TEVA | 2022-05-04 12:28 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.35 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.99 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:25 | GW012-Mannul-SP30-MD-01 | 1.009 | 0.047 | 408.9 | 5.5 | 133.5 | 0.079 | -11.3 | -1.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -11.4 | -2 | 23.81 | 21.96 | 1.084 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 76.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.63 | TEVA | 2022-05-04 12:25 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.43 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.4 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:22 | GW012-Mannul-SP30-MD-01 | 1.099 | 0.066 | 381.1 | 5.1 | 317.4 | 0.202 | -9.6 | -3.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -9.4 | -2.9 | 23.78 | 20.45 | 1.162 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 282 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 85.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.8 | TEVA | 2022-05-04 12:23 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.76 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.68 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:20 | GW012-Mannul-SP30-MD-01 | 1.117 | 0.064 | 427.8 | 5.8 | 275.9 | 0.249 | -11.5 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -11.5 | -0.8 | 25.66 | 21.51 | 1.193 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 414 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 114.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.37 | TEVA | 2022-05-04 12:20 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.82 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.06 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:17 | GW012-Mannul-SP30-MD-01 | 1.098 | 0.041 | 448.5 | 6.1 | 245.6 | 0.241 | -11.4 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -11.5 | -1.5 | 25.87 | 22.32 | 1.159 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 414 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 122.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.72 | TEVA | 2022-05-04 12:17 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.52 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.31 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 80 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:14 | GW012-Mannul-SP30-MD-01 | 1.094 | 0.042 | 437.7 | 5.9 | 109.3 | 0.113 | -11.1 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -11.1 | -1.7 | 25.37 | 22.12 | 1.147 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 70.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.4 | TEVA | 2022-05-04 12:15 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.31 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.37 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:12 | GW012-Mannul-SP30-MD-01 | 1.047 | 0.042 | 305.7 | 4.1 | 199.5 | 0.048 | -13.8 | -6.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -13.8 | -6.5 | 20.84 | 18.6 | 1.121 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 424 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 135.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 19.9 | TEVA | 2022-05-04 12:12 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 19.02 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 62.01 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 70 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:09 | GW012-Mannul-SP30-MD-01 | 1.133 | 0.061 | 471.2 | 6.4 | 220.3 | 0.281 | -12.5 | -2.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -12.7 | -2.2 | 26.63 | 22.59 | 1.179 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 288 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 44 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.77 | TEVA | 2022-05-04 12:10 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.73 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.77 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:07 | GW012-Mannul-SP30-MD-01 | 1.141 | 0.047 | 384.5 | 5.2 | 29.1 | 0.048 | -10.9 | -2.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -10.8 | -2.6 | 24.39 | 20.47 | 1.191 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 344 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 104.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.33 | TEVA | 2022-05-04 12:07 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.45 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.22 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:04 | GW012-Mannul-SP30-MD-01 | 1.107 | 0.047 | 411.6 | 5.6 | 346.3 | 0.103 | -12 | -2.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -11.9 | -2.5 | 24.58 | 21.4 | 1.149 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 394 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 50.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.76 | TEVA | 2022-05-04 12:04 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.47 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.84 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 12:01 | GW012-Mannul-SP30-MD-01 | 1.154 | 0.045 | 304.1 | 4.1 | 14.1 | 0.031 | -15 | -5.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.9 | MDx |  |  |  |  | -15 | -5.7 | 21.64 | 17.91 | 1.208 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 416 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 130.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 20.81 | TEVA | 2022-05-04 12:02 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 18.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 62.02 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 58 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:59 | GW012-Mannul-SP30-MD-01 | 1.164 | 0.049 | 387.7 | 5.2 | 285.6 | 0.099 | -13.5 | -5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.9 | MDx |  |  |  |  | -13.4 | -4.9 | 24.55 | 19.89 | 1.234 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 498 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 164.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.7 | TEVA | 2022-05-04 11:59 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.36 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.42 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:56 | GW012-Mannul-SP30-MD-01 | 1.102 | 0.05 | 357 | 4.8 | 26.7 | 0.072 | -11.7 | -4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -11.7 | -4 | 22.67 | 19.78 | 1.146 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 418 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 116.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.06 | TEVA | 2022-05-04 11:57 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.02 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 68.21 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 106 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:54 | GW012-Mannul-SP30-MD-01 | 1.061 | 0.055 | 391.5 | 5.3 | 251.4 | 0.066 | -12.2 | -2.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.2 | -2.1 | 23.34 | 20.97 | 1.113 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 450 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 35.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.7 | TEVA | 2022-05-04 11:54 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.4 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.85 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 19 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:51 | GW012-Mannul-SP30-MD-01 | 1.115 | 0.077 | 465.7 | 6.3 | 205.1 | 0.202 | -11.3 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -11.5 | -1 | 27.25 | 22.74 | 1.198 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 426 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 172.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.47 | TEVA | 2022-05-04 11:51 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.85 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.96 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:49 | GW012-Mannul-SP30-MD-01 | 1.067 | 0.041 | 372 | 5 | 335 | 0.175 | -11 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -10.9 | -1.4 | 23.22 | 20.54 | 1.131 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 402 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 56.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.17 | TEVA | 2022-05-04 11:49 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.77 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.43 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:46 | GW012-Mannul-SP30-MD-01 | 1.017 | 0.055 | 372.7 | 5 | 202.4 | 0.129 | -12.3 | -3.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -12.5 | -3.5 | 22.62 | 20.89 | 1.083 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 374 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 56.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 21.68 | TEVA | 2022-05-04 11:46 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.32 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.38 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:43 | GW012-Mannul-SP30-MD-01 | 1.067 | 0.057 | 411.5 | 5.6 | 278.3 | 0.153 | -11 | -1.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.9 | MDx |  |  |  |  | -11 | -1.7 | 24.57 | 21.54 | 1.14 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 113.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.35 | TEVA | 2022-05-04 11:44 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.89 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.73 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 102 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:41 | GW012-Mannul-SP30-MD-01 | 1.027 | 0.051 | 433.7 | 5.9 | 292.3 | 0.058 | -10.6 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -10.6 | -1.1 | 24.8 | 22.57 | 1.099 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 96.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.55 | TEVA | 2022-05-04 11:41 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.93 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.48 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:38 | GW012-Mannul-SP30-MD-01 | 1.148 | 0.045 | 376.1 | 5.1 | 205.8 | 0.016 | -9.5 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -9.5 | -1.7 | 24.38 | 20 | 1.219 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 436 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 32.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.14 | TEVA | 2022-05-04 11:38 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.16 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.14 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 116 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:34 | GW012-Mannul-SP30-MD-01 | 1.076 | 0.041 | 412.1 | 5.6 | 266.2 | 0.09 | -12.4 | -2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.4 | -1.9 | 24.21 | 21.58 | 1.122 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 394 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 173.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.45 | TEVA | 2022-05-04 11:34 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.79 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.72 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:31 | GW012-Mannul-SP30-MD-01 | 1.104 | 0.055 | 289.8 | 3.9 | 153.3 | 0.036 | -9.5 | 5.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -9.5 | 5.4 | 20.82 | 18.06 | 1.153 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 498 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 33.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 19.85 | TEVA | 2022-05-04 11:32 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 17.99 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 60.13 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:29 | GW012-Mannul-SP30-MD-01 | 1.057 | 0.051 | 434 | 5.9 | 207.2 | 0.367 | -11.5 | -2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -11.8 | -1.8 | 25.37 | 22.11 | 1.147 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 344 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 127.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.87 | TEVA | 2022-05-04 11:29 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.58 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.66 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:26 | GW012-Mannul-SP30-MD-01 | 1.02 | 0.037 | 346.5 | 4.7 | 264.3 | 0.082 | -12.1 | -3.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.1 | -3 | 21.74 | 19.96 | 1.089 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 380 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 52.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 20.92 | TEVA | 2022-05-04 11:26 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.51 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 66.75 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:23 | GW012-Mannul-SP30-MD-01 | 1.048 | 0.048 | 297.1 | 4 | 115.9 | 0.073 | -11.5 | -2.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -11.6 | -2.9 | 20.45 | 18.3 | 1.118 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 498 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 34.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 19.61 | TEVA | 2022-05-04 11:24 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 18.71 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 61.85 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 152 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:21 | GW012-Mannul-SP30-MD-01 | 1.086 | 0.06 | 407.8 | 5.5 | 293.2 | 0.057 | -10.9 | -2.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -10.9 | -2.1 | 24.05 | 21.07 | 1.141 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 416 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 14.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.45 | TEVA | 2022-05-04 11:21 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.59 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.03 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 11:18 | GW012-Mannul-SP30-MD-01 | 1.04 | 0.059 | 431.1 | 5.8 | 300.2 | 0.166 | -10.5 | -2.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -10.4 | -2 | 24.6 | 22.29 | 1.104 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 364 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 146.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.58 | TEVA | 2022-05-04 11:18 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.68 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.66 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 56 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:29 | GW012-Mannul-SP30-MD-01 | 1.029 | 0.054 | 389.2 | 5.3 | 291.9 | 0.075 | -10.5 | -3.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -10.5 | -3.4 | 23.13 | 21.23 | 1.09 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 274 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 74.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.3 | TEVA | 2022-05-04 10:29 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.67 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.07 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:27 | GW012-Mannul-SP30-MD-01 | 1.072 | 0.08 | 463.4 | 6.3 | 225.7 | 0.069 | -11.9 | -2.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -11.9 | -2.4 | 25.72 | 22.65 | 1.136 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 312 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 17.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.91 | TEVA | 2022-05-04 10:27 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.24 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.66 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 18 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:24 | GW012-Mannul-SP30-MD-01 | 1.044 | 0.036 | 445.5 | 6 | 242.2 | 0.103 | -12 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -12.1 | -1.6 | 25.11 | 22.8 | 1.101 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 408 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 8.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.05 | TEVA | 2022-05-04 10:24 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.03 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.85 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:21 | GW012-Mannul-SP30-MD-01 | 1.05 | 0.045 | 372.7 | 5 | 235.3 | 0.086 | -10.7 | -2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -10.8 | -2 | 22.93 | 20.84 | 1.1 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 12.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.04 | TEVA | 2022-05-04 10:22 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.17 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 72 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:19 | GW012-Mannul-SP30-MD-01 | 1.033 | 0.045 | 449.1 | 6.1 | 220.9 | 0.105 | -10.8 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -10.9 | -1.4 | 25.09 | 22.95 | 1.094 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 402 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 84.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.02 | TEVA | 2022-05-04 10:19 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.26 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.27 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:16 | GW012-Mannul-SP30-MD-01 | 1.115 | 0.037 | 387.1 | 5.2 | 272.4 | 0.079 | -11.7 | -2.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -11.7 | -2.3 | 23.89 | 20.55 | 1.162 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 81 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.12 | TEVA | 2022-05-04 10:16 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.73 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.3 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:14 | GW012-Mannul-SP30-MD-01 | 1.131 | 0.064 | 392.4 | 5.3 | 168.7 | 0.202 | -13.3 | -3.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -13.5 | -3.1 | 24.3 | 20.12 | 1.208 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 376 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 8.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.48 | TEVA | 2022-05-04 10:14 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.75 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.29 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:11 | GW012-Mannul-SP30-MD-01 | 1.088 | 0.065 | 424.9 | 5.7 | 201.6 | 0.218 | -12.9 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -13.1 | -1.5 | 24.82 | 21.53 | 1.153 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 422 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 27.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.97 | TEVA | 2022-05-04 10:11 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.47 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:09 | GW012-Mannul-SP30-MD-01 | 1.188 | 0.053 | 392.1 | 5.3 | 212.2 | 0.268 | -12.5 | -2.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -12.8 | -2.2 | 25.26 | 20.47 | 1.234 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 340 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 170.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.07 | TEVA | 2022-05-04 10:09 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.26 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.91 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:06 | GW012-Mannul-SP30-MD-01 | 1.187 | 0.044 | 431.3 | 5.8 | 230 | 0.448 | -13.2 | -3.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -13.5 | -2.8 | 26.67 | 21.19 | 1.258 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 256 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 114.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.21 | TEVA | 2022-05-04 10:06 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.25 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.59 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:03 | GW012-Mannul-SP30-MD-01 | 1.087 | 0.045 | 383.5 | 5.2 | 294.3 | 0.057 | -11.5 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -11.5 | -1.6 | 23.91 | 20.84 | 1.147 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 446 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 152.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.7 | TEVA | 2022-05-04 10:04 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.89 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.66 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 10:01 | GW012-Mannul-SP30-MD-01 | 1.125 | 0.041 | 401.5 | 5.4 | 213.7 | 0.095 | -12.1 | -2.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -12.2 | -2.8 | 24.28 | 20.71 | 1.172 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 344 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 155.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.66 | TEVA | 2022-05-04 10:01 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.51 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:58 | GW012-Mannul-SP30-MD-01 | 1.091 | 0.044 | 458.7 | 6.2 | 237.3 | 0.213 | -12.3 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -12.4 | -1 | 26.41 | 22.57 | 1.17 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 430 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 143.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.93 | TEVA | 2022-05-04 09:58 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.85 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.98 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 102 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:55 | GW012-Mannul-SP30-MD-01 | 1.144 | 0.034 | 509.4 | 6.9 | 208.7 | 0.28 | -11.6 | -1.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -11.8 | -1.3 | 27.57 | 23.61 | 1.168 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 314 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 137.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.93 | TEVA | 2022-05-04 09:55 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.54 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.67 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:52 | GW012-Mannul-SP30-MD-01 | 1.205 | 0.044 | 476.5 | 6.4 | 104.7 | 0.025 | -11.5 | -1.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -11.5 | -1.9 | 27.65 | 22.22 | 1.244 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 370 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 177.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.72 | TEVA | 2022-05-04 09:52 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.18 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.56 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 102 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:49 | GW012-Mannul-SP30-MD-01 | 1.029 | 0.046 | 452.2 | 6.1 | 203 | 0.145 | -10.7 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -10.8 | -1.6 | 25.33 | 22.67 | 1.117 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 396 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 163.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.08 | TEVA | 2022-05-04 09:50 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.41 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.04 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:47 | GW012-Mannul-SP30-MD-01 | 1.088 | 0.038 | 435 | 5.9 | 276.2 | 0.087 | -11.7 | -2.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -11.7 | -2 | 25.43 | 21.88 | 1.163 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 396 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 57.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.26 | TEVA | 2022-05-04 09:47 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.3 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.64 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 17 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:44 | GW012-Mannul-SP30-MD-01 | 1.125 | 0.04 | 426.2 | 5.8 | 307.3 | 0.152 | -11.5 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -11.4 | -1.6 | 25.59 | 21.27 | 1.204 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 384 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 87.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.41 | TEVA | 2022-05-04 09:44 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.68 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.62 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:27 | GW012-Mannul-SP30-MD-01 | 1.075 | 0.048 | 385.1 | 5.2 | 61.7 | 0.028 | -12.1 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.1 | -1.3 | 23.82 | 20.76 | 1.147 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 390 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 143.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.66 | TEVA | 2022-05-04 09:28 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.08 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.87 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:24 | GW012-Mannul-SP30-MD-01 | 1.113 | 0.04 | 443.5 | 6 | 184.5 | 0.152 | -11.8 | -1.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12 | -1.9 | 25.49 | 21.8 | 1.169 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 400 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 73.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.77 | TEVA | 2022-05-04 09:25 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.26 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.82 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:22 | GW012-Mannul-SP30-MD-01 | 1.161 | 0.054 | 440.2 | 5.9 | 139.1 | 0.13 | -10.7 | -2.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -10.8 | -2.7 | 26.24 | 20.91 | 1.255 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 344 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 141.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.2 | TEVA | 2022-05-04 09:22 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.71 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.64 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:19 | GW012-Mannul-SP30-MD-01 | 1.041 | 0.052 | 430.9 | 5.8 | 247.8 | 0.139 | -10.4 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -10.4 | -0.8 | 24.63 | 22.26 | 1.107 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 376 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 112 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.58 | TEVA | 2022-05-04 09:19 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.65 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.61 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:16 | GW012-Mannul-SP30-MD-01 | 1.083 | 0.075 | 456.3 | 6.2 | 249.6 | 0.147 | -12.2 | -2.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -12.3 | -2.1 | 25.49 | 22.57 | 1.129 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 33.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.81 | TEVA | 2022-05-04 09:16 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.91 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.27 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 100 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:14 | GW012-Mannul-SP30-MD-01 | 1.044 | 0.066 | 431.9 | 5.8 | 260.8 | 0.156 | -11.8 | -2.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.4 | MDx |  |  |  |  | -11.9 | -2.2 | 25.26 | 22.12 | 1.142 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 368 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 155.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.68 | TEVA | 2022-05-04 09:14 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.68 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.94 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:11 | GW012-Mannul-SP30-MD-01 | 1.094 | 0.047 | 415.7 | 5.6 | 76.5 | 0.035 | -12.6 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.3 | MDx |  |  |  |  | -12.6 | -1.4 | 24.29 | 21.25 | 1.143 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 382 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 84.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.74 | TEVA | 2022-05-04 09:11 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.69 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.76 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:07 | GW012-Mannul-SP30-MD-01 | 1.023 | 0.08 | 519.1 | 7 | 232.6 | 0.15 | -11.4 | -0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.3 | MDx |  |  |  |  | -11.5 | -0.5 | 27.31 | 24.48 | 1.116 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 430 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 135.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.73 | TEVA | 2022-05-04 09:07 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 25.15 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 82.11 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:05 | GW012-Mannul-SP30-MD-01 | 1.018 | 0.044 | 508.1 | 6.9 | 226.3 | 0.201 | -12.1 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.1 | MDx |  |  |  |  | -12.2 | -0.7 | 26.7 | 24.16 | 1.105 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 418 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 123 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.4 | TEVA | 2022-05-04 09:05 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.94 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 81.07 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 82 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 09:02 | GW012-Mannul-SP30-MD-01 | 1.111 | 0.053 | 454.5 | 6.1 | 240.6 | 0.069 | -10.2 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36 | MDx |  |  |  |  | -10.2 | -1.3 | 26.03 | 22.26 | 1.169 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 330 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 6.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.05 | TEVA | 2022-05-04 09:02 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.54 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.08 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 70 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 08:59 | GW012-Mannul-SP30-MD-01 | 1.101 | 0.056 | 537.5 | 7.3 | 270.9 | 0.326 | -13.2 | -0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36 | MDx |  |  |  |  | -13.2 | -0.2 | 28.26 | 24.16 | 1.17 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 338 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 109 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 27.15 | TEVA | 2022-05-04 09:00 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.66 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 83.52 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 76 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 08:57 | GW012-Mannul-SP30-MD-01 | 1.076 | 0.06 | 460.7 | 6.2 | 237 | 0.087 | -11.4 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 35.8 | MDx |  |  |  |  | -11.5 | -1.5 | 26.02 | 23.14 | 1.125 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 366 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 142.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.8 | TEVA | 2022-05-04 08:57 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.05 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.72 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 70 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 08:54 | GW012-Mannul-SP30-MD-01 | 1.126 | 0.045 | 434.6 | 5.9 | 291.8 | 0.129 | -11.2 | -2.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 35.8 | MDx |  |  |  |  | -11.1 | -2 | 25.25 | 21.86 | 1.155 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 296 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 87.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.64 | TEVA | 2022-05-04 08:54 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.89 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.51 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 08:51 | GW012-Mannul-SP30-MD-01 | 1.115 | 0.052 | 476.8 | 6.4 | 231.5 | 0.184 | -11.5 | -0.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 35.4 | MDx |  |  |  |  | -11.7 | -0.4 | 26.58 | 22.44 | 1.185 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 400 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 149.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.71 | TEVA | 2022-05-04 08:52 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.06 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.72 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 16 | 120 | cfpharmtech\mi1chen | 2022-05-04 08:48 | GW012-Mannul-SP30-MD-01 | 1.066 | 0.035 | 510.1 | 6.9 | 193.9 | 0.221 | -12.7 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 35.3 | MDx |  |  |  |  | -12.9 | -0.9 | 27.07 | 24.07 | 1.124 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 408 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 139.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26 | TEVA | 2022-05-04 08:48 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.39 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 81.42 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 16:28 | GW012-Mannul-SP30-MD-01 | 1.091 | 0.044 | 378.7 | 5.1 | 232.1 | 0.242 | -11.6 | 5.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -11.8 | 5.7 | 24.24 | 20.34 | 1.191 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 404 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 131.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.65 | TEVA | 2022-04-29 16:28 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.77 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.4 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 16:26 | GW012-Mannul-SP30-MD-01 | 1.008 | 0.066 | 388.2 | 5.2 | 219.4 | 0.176 | -9.2 | -2.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -9.3 | -2.7 | 23.33 | 21.3 | 1.095 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 306 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 111.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.03 | TEVA | 2022-04-29 16:26 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.86 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.06 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 72 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 16:23 | GW012-Mannul-SP30-MD-01 | 1.076 | 0.039 | 418.4 | 5.6 | 318.8 | 0.171 | -12.8 | -2.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -12.6 | -2.4 | 24.99 | 21.67 | 1.153 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 336 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 162 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.63 | TEVA | 2022-04-29 16:23 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.97 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.22 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 16:20 | GW012-Mannul-SP30-MD-01 | 1.041 | 0.034 | 398.4 | 5.4 | 180.9 | 0.117 | -11.6 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -11.7 | -1 | 23.67 | 21.14 | 1.12 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 178.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.69 | TEVA | 2022-04-29 16:21 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.79 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.46 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 16:17 | GW012-Mannul-SP30-MD-01 | 1.099 | 0.083 | 453.4 | 6.1 | 216.9 | 0.127 | -9.7 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -9.8 | -0.8 | 25.52 | 22.04 | 1.158 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 404 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 23 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.89 | TEVA | 2022-04-29 16:18 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.66 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.88 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 16:15 | GW012-Mannul-SP30-MD-01 | 1.143 | 0.036 | 408.3 | 5.5 | 202.7 | 0.068 | -12.9 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -13 | -0.8 | 24.46 | 20.99 | 1.165 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 396 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 178.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.04 | TEVA | 2022-04-29 16:15 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.07 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 110 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 16:12 | GW012-Mannul-SP30-MD-01 | 1.039 | 0.056 | 424.6 | 5.7 | 186.6 | 0.204 | -8.9 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -9.1 | -1 | 24.79 | 22.34 | 1.11 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 132.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.38 | TEVA | 2022-04-29 16:12 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.51 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.13 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 108 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 16:09 | GW012-Mannul-SP30-MD-01 | 1.084 | 0.052 | 486.1 | 6.5 | 253.2 | 0.285 | -13 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -13 | -0.8 | 27.24 | 23.21 | 1.174 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 410 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 106.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.59 | TEVA | 2022-04-29 16:09 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.61 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.71 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 128 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 16:06 | GW012-Mannul-SP30-MD-01 | 1.02 | 0.056 | 434.5 | 5.9 | 122.4 | 0.171 | -12.8 | -1.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -12.9 | -2 | 24.82 | 22.28 | 1.114 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 346 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 27.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.5 | TEVA | 2022-04-29 16:07 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.03 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.57 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 16:04 | GW012-Mannul-SP30-MD-01 | 1.091 | 0.051 | 404.2 | 5.4 | 332.3 | 0.122 | -10.6 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -10.5 | -1.4 | 24 | 21.37 | 1.123 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 410 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 40.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.4 | TEVA | 2022-04-29 16:04 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.44 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.47 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:59 | GW012-Mannul-SP30-MD-01 | 1.18 | 0.075 | 457.4 | 6.2 | 244.8 | 0.225 | -10.6 | -2.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.9 | MDx |  |  |  |  | -10.7 | -2.1 | 26.08 | 21.42 | 1.218 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 314 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 1.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.89 | TEVA | 2022-04-29 15:59 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.94 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.61 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:56 | GW012-Mannul-SP30-MD-01 | 1.087 | 0.07 | 406.7 | 5.5 | 280 | 0.251 | -11.1 | -0.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -11.1 | -0.2 | 24.35 | 21.37 | 1.139 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 498 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 100.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.47 | TEVA | 2022-04-29 15:56 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.59 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.38 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:54 | GW012-Mannul-SP30-MD-01 | 1.137 | 0.06 | 487.8 | 6.6 | 215 | 0.137 | -12.3 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.4 | -1.2 | 28.07 | 22.8 | 1.231 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 360 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 156.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.26 | TEVA | 2022-04-29 15:54 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.09 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.88 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:50 | GW012-Mannul-SP30-MD-01 | 1.129 | 0.064 | 457.2 | 6.2 | 225.6 | 0.211 | -11.8 | -1.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12 | -1.8 | 25.81 | 22.24 | 1.16 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 360 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 2.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.31 | TEVA | 2022-04-29 15:50 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.42 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.64 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 15 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:47 | GW012-Mannul-SP30-MD-01 | 1.059 | 0.055 | 408.1 | 5.5 | 229.6 | 0.138 | -10.3 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -10.3 | -0.9 | 24.24 | 21.66 | 1.119 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 410 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 81.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.22 | TEVA | 2022-04-29 15:47 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.92 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.27 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:44 | GW012-Mannul-SP30-MD-01 | 1.11 | 0.074 | 425.6 | 5.7 | 208.6 | 0.314 | -12 | -2.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.3 | -2.3 | 25.06 | 21.27 | 1.178 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 370 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 5.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.2 | TEVA | 2022-04-29 15:44 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.81 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.54 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:41 | GW012-Mannul-SP30-MD-01 | 1.083 | 0.062 | 428.8 | 5.8 | 288.7 | 0.092 | -12.5 | -0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -12.5 | -0.7 | 25.9 | 21.67 | 1.195 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 378 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 122.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.02 | TEVA | 2022-04-29 15:41 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.18 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.36 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:37 | GW012-Mannul-SP30-MD-01 | 1.036 | 0.05 | 432.9 | 5.8 | 247.5 | 0.12 | -11.1 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -11.1 | -1.1 | 25.6 | 22.25 | 1.15 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 404 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 173.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.61 | TEVA | 2022-04-29 15:38 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.79 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.48 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:34 | GW012-Mannul-SP30-MD-01 | 1.143 | 0.045 | 446.5 | 6 | 309.1 | 0.063 | -10.8 | -2.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -10.7 | -2.2 | 25.68 | 21.54 | 1.192 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 340 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 21.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.17 | TEVA | 2022-04-29 15:35 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.02 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.85 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:31 | GW012-Mannul-SP30-MD-01 | 1.103 | 0.044 | 384.8 | 5.2 | 282.9 | 0.22 | -10.4 | -2.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -10.3 | -2.2 | 23.84 | 20.33 | 1.173 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 354 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 70.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.91 | TEVA | 2022-04-29 15:31 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.78 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.26 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 56 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:29 | GW012-Mannul-SP30-MD-01 | 1.108 | 0.058 | 389.6 | 5.2 | 247.9 | 0.063 | -12.1 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -12.1 | -1.5 | 23.9 | 20.56 | 1.163 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 424 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 87.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.15 | TEVA | 2022-04-29 15:29 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.89 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.7 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 78 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:26 | GW012-Mannul-SP30-MD-01 | 1.089 | 0.042 | 379.3 | 5.1 | 212.4 | 0.076 | -11.4 | -0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -11.4 | -0.6 | 23.41 | 20.86 | 1.122 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 432 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 111.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.64 | TEVA | 2022-04-29 15:26 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.79 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.51 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:23 | GW012-Mannul-SP30-MD-01 | 1.014 | 0.042 | 412.9 | 5.6 | 261.7 | 0.121 | -10.3 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -10.3 | -0.9 | 23.87 | 22.14 | 1.078 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 384 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 2.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.77 | TEVA | 2022-04-29 15:23 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.46 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.92 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 98 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:21 | GW012-Mannul-SP30-MD-01 | 1.172 | 0.044 | 400 | 5.4 | 263.5 | 0.154 | -10.1 | -2.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -10.1 | -2.6 | 25.29 | 20.53 | 1.232 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 300 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 99.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.13 | TEVA | 2022-04-29 15:21 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.59 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.66 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:18 | GW012-Mannul-SP30-MD-01 | 1.049 | 0.046 | 401.2 | 5.4 | 94.6 | 0.028 | -9.9 | -1.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -9.9 | -1.9 | 24.16 | 21.37 | 1.13 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 426 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 79.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.87 | TEVA | 2022-04-29 15:18 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.8 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.52 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 82 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:15 | GW012-Mannul-SP30-MD-01 | 1.032 | 0.068 | 442.4 | 6 | 263.5 | 0.244 | -12.4 | -2.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -12.5 | -1.9 | 25.48 | 22.3 | 1.143 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 324 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 89.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.83 | TEVA | 2022-04-29 15:15 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.1 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.84 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:12 | GW012-Mannul-SP30-MD-01 | 1.012 | 0.054 | 494.1 | 6.7 | 266.9 | 0.208 | -11.7 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -11.7 | -0.9 | 25.74 | 24.18 | 1.065 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 426 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 48.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.95 | TEVA | 2022-04-29 15:12 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.66 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 79.28 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 110 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:09 | GW012-Mannul-SP30-MD-01 | 1.042 | 0.071 | 435.6 | 5.9 | 213.7 | 0.156 | -11.7 | -2.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.8 | -2.3 | 24.99 | 22.37 | 1.117 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 354 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 172.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.77 | TEVA | 2022-04-29 15:09 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.82 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.76 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:06 | GW012-Mannul-SP30-MD-01 | 1.174 | 0.05 | 461.3 | 6.2 | 108.6 | 0.047 | -11.7 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.7 | -1.2 | 26.84 | 21.39 | 1.255 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 168.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.95 | TEVA | 2022-04-29 15:06 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.1 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.19 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 14 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:04 | GW012-Mannul-SP30-MD-01 | 1.033 | 0.034 | 471.4 | 6.4 | 245.6 | 0.124 | -11.3 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.4 | -0.8 | 25.77 | 23.36 | 1.103 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 414 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 120.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.59 | TEVA | 2022-04-29 15:04 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.8 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.56 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 15:00 | GW012-Mannul-SP30-MD-01 | 1.008 | 0.047 | 428.3 | 5.8 | 131.9 | 0.098 | -11 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -11.1 | -1.2 | 24.31 | 22.35 | 1.088 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 394 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 21.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.13 | TEVA | 2022-04-29 15:00 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.95 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.18 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:57 | GW012-Mannul-SP30-MD-01 | 1.197 | 0.098 | 489.5 | 6.6 | 336.5 | 0.092 | -12.4 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12.3 | -1.4 | 27.01 | 21.48 | 1.258 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 378 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 31 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 27.1 | TEVA | 2022-04-29 14:58 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.65 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 81.6 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:52 | GW012-Mannul-SP30-MD-01 | 1.037 | 0.057 | 459.8 | 6.2 | 140.7 | 0.063 | -12.1 | -1.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.8 | MDx |  |  |  |  | -12.2 | -1.9 | 25.24 | 22.29 | 1.133 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 332 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 42.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.35 | TEVA | 2022-04-29 14:52 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.49 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.02 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:47 | GW012-Mannul-SP30-MD-01 | 1.039 | 0.046 | 436.5 | 5.9 | 236 | 0.02 | -10.9 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -10.9 | -1.1 | 24.62 | 22.24 | 1.107 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 408 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 80.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.72 | TEVA | 2022-04-29 14:48 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.84 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.95 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:43 | GW012-Mannul-SP30-MD-01 | 1.132 | 0.046 | 488.6 | 6.6 | 228.2 | 0.195 | -11.4 | -0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.3 | MDx |  |  |  |  | -11.5 | -0.5 | 26.83 | 22.59 | 1.188 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 440 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 172.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.22 | TEVA | 2022-04-29 14:43 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.16 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 79.72 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:40 | GW012-Mannul-SP30-MD-01 | 1.087 | 0.052 | 489.9 | 6.6 | 263.4 | 0.191 | -10.8 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.3 | MDx |  |  |  |  | -10.8 | -0.9 | 26.47 | 22.94 | 1.154 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 354 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 107.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.73 | TEVA | 2022-04-29 14:40 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.66 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.36 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:36 | GW012-Mannul-SP30-MD-01 | 1.154 | 0.046 | 411.5 | 5.5 | 272.4 | 0.144 | -11.4 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.4 | MDx |  |  |  |  | -11.4 | -1.4 | 25.54 | 20.47 | 1.248 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 382 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 119.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.29 | TEVA | 2022-04-29 14:36 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.05 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.13 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:33 | GW012-Mannul-SP30-MD-01 | 1.09 | 0.047 | 427.3 | 5.8 | 209.4 | 0.134 | -11.9 | -1.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.7 | MDx |  |  |  |  | -12 | -1.8 | 24.69 | 21.85 | 1.13 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 354 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 91.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.03 | TEVA | 2022-04-29 14:34 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.53 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:30 | GW012-Mannul-SP30-MD-01 | 1.145 | 0.049 | 477.2 | 6.4 | 245.7 | 0.057 | -12.1 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.9 | MDx |  |  |  |  | -12.2 | -1.5 | 26.49 | 22.58 | 1.173 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 390 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 21.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.1 | TEVA | 2022-04-29 14:30 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.79 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.72 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 98 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:28 | GW012-Mannul-SP30-MD-01 | 1.028 | 0.042 | 501.3 | 6.8 | 236.5 | 0.206 | -11.7 | -1.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39 | MDx |  |  |  |  | -11.8 | -1.2 | 26.52 | 24.27 | 1.093 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 350 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 160.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.33 | TEVA | 2022-04-29 14:28 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.63 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.14 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 70 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:25 | GW012-Mannul-SP30-MD-01 | 1.039 | 0.043 | 440.5 | 5.9 | 229.4 | 0.07 | -9.8 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.2 | MDx |  |  |  |  | -9.9 | -1.7 | 24.65 | 22.7 | 1.086 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 370 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 10.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.83 | TEVA | 2022-04-29 14:25 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.94 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.78 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:22 | GW012-Mannul-SP30-MD-01 | 1.02 | 0.051 | 390.6 | 5.3 | 350.1 | 0.114 | -8.4 | -4.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.5 | MDx |  |  |  |  | -8.3 | -4.6 | 23.42 | 21.07 | 1.111 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 274 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 164 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.26 | TEVA | 2022-04-29 14:22 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.82 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.4 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:20 | GW012-Mannul-SP30-MD-01 | 1.035 | 0.046 | 511.1 | 6.9 | 222.5 | 0.053 | -13 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.6 | MDx |  |  |  |  | -13 | -1.2 | 27.06 | 24.25 | 1.116 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 396 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 140.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.68 | TEVA | 2022-04-29 14:20 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.81 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.81 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 82 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:16 | GW012-Mannul-SP30-MD-01 | 1.052 | 0.063 | 504 | 6.8 | 241.8 | 0.213 | -11.4 | -2.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.9 | MDx |  |  |  |  | -11.5 | -2.1 | 27.58 | 23.31 | 1.183 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 402 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 168 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.7 | TEVA | 2022-04-29 14:17 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.42 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.74 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 13 | 120 | cfpharmtech\mi1chen | 2022-04-29 14:13 | GW012-Mannul-SP30-MD-01 | 1.028 | 0.059 | 445.8 | 6 | 351.2 | 0.064 | -10.3 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 40 | MDx |  |  |  |  | -10.3 | -1.5 | 24.69 | 22.97 | 1.075 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 414 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 21.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.84 | TEVA | 2022-04-29 14:13 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.21 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.53 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 15:02 | GW012-Mannul-SP30-MD-01 | 1.035 | 0.052 | 449.3 | 6.1 | 358.4 | 0.121 | -11.3 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -11.2 | -1.1 | 25.2 | 23.01 | 1.095 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 380 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 3.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.01 | TEVA | 2022-04-28 15:02 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.2 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.66 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 98 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:59 | GW012-Mannul-SP30-MD-01 | 1.163 | 0.071 | 480.4 | 6.5 | 162.3 | 0.131 | -13.1 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -13.2 | -1 | 27.23 | 22.05 | 1.235 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 153.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.38 | TEVA | 2022-04-28 14:59 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.69 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 79.53 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:56 | GW012-Mannul-SP30-MD-01 | 1.17 | 0.036 | 405.8 | 5.5 | 342.2 | 0.142 | -9.6 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -9.5 | -1.4 | 25.21 | 20.26 | 1.244 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 418 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 114 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.26 | TEVA | 2022-04-28 14:56 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.74 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.46 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:53 | GW012-Mannul-SP30-MD-01 | 1.063 | 0.03 | 443.4 | 6 | 183.4 | 0.117 | -11.7 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.8 | -1 | 25.23 | 22.29 | 1.132 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 396 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 157.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.19 | TEVA | 2022-04-28 14:53 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.75 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.09 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:50 | GW012-Mannul-SP30-MD-01 | 1.059 | 0.051 | 474.6 | 6.4 | 290 | 0.074 | -11.5 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.4 | -0.9 | 25.66 | 22.92 | 1.12 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 388 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 20.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.99 | TEVA | 2022-04-28 14:51 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.59 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.6 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:48 | GW012-Mannul-SP30-MD-01 | 1.178 | 0.067 | 525.5 | 7.1 | 214.4 | 0.092 | -13.5 | -0.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -13.6 | -0.3 | 28.65 | 22.58 | 1.268 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 410 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 6.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 27.74 | TEVA | 2022-04-28 14:48 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.54 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 84.61 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:45 | GW012-Mannul-SP30-MD-01 | 1.029 | 0.048 | 473.9 | 6.4 | 274.9 | 0.299 | -12.1 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -12.1 | -0.5 | 25.5 | 23.49 | 1.086 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 328 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 28.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.6 | TEVA | 2022-04-28 14:45 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.92 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.86 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:43 | GW012-Mannul-SP30-MD-01 | 1.046 | 0.06 | 418.3 | 5.7 | 198.7 | 0.076 | -13.7 | -5.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -13.8 | -5.3 | 24.47 | 21.67 | 1.13 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 284 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 26.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.3 | TEVA | 2022-04-28 14:43 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.27 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.43 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:40 | GW012-Mannul-SP30-MD-01 | 1.166 | 0.064 | 456.8 | 6.2 | 232.6 | 0.199 | -13.3 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -13.4 | -1 | 26.4 | 20.91 | 1.263 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 352 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 140.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.77 | TEVA | 2022-04-28 14:40 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.11 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.13 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:37 | GW012-Mannul-SP30-MD-01 | 1.04 | 0.048 | 419.5 | 5.7 | 244.6 | 0.023 | -13.1 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -13.2 | -0.8 | 24.59 | 22.13 | 1.111 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 352 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 74.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.24 | TEVA | 2022-04-28 14:37 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.35 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.97 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:33 | GW012-Mannul-SP30-MD-01 | 1.106 | 0.066 | 427.6 | 5.8 | 216.1 | 0.238 | -12.3 | -1.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -12.5 | -1.2 | 25.42 | 21.81 | 1.166 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 370 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 24.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.21 | TEVA | 2022-04-28 14:33 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.9 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.34 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:30 | GW012-Mannul-SP30-MD-01 | 1.115 | 0.039 | 496.8 | 6.7 | 233.5 | 0.14 | -12.4 | -0.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12.5 | -0.1 | 27.16 | 23.35 | 1.164 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 378 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 22.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.26 | TEVA | 2022-04-28 14:30 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.57 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.19 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 98 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:27 | GW012-Mannul-SP30-MD-01 | 1.098 | 0.047 | 458 | 6.2 | 224.5 | 0.103 | -11.9 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12 | -1.1 | 25.6 | 22.29 | 1.149 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 346 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 140.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.02 | TEVA | 2022-04-28 14:28 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.79 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.06 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:25 | GW012-Mannul-SP30-MD-01 | 1.033 | 0.068 | 461.2 | 6.3 | 216.6 | 0.057 | -12.4 | -0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12.5 | -0.6 | 25.53 | 22.69 | 1.125 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 404 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 27.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.39 | TEVA | 2022-04-28 14:25 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.62 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.08 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 100 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 12 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:22 | GW012-Mannul-SP30-MD-01 | 1.144 | 0.068 | 455.1 | 6.2 | 140.5 | 0.071 | -12.8 | -0.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12.8 | -0.4 | 25.88 | 21.75 | 1.19 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 164.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.46 | TEVA | 2022-04-28 14:22 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.26 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.96 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:19 | GW012-Mannul-SP30-MD-01 | 1.056 | 0.044 | 439.4 | 6 | 294.3 | 0.074 | -12.2 | -0.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12.2 | -0.3 | 24.95 | 22.25 | 1.121 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 400 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 128.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.03 | TEVA | 2022-04-28 14:20 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.75 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.71 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:17 | GW012-Mannul-SP30-MD-01 | 1.052 | 0.063 | 490.6 | 6.7 | 152.3 | 0.198 | -12.8 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -13 | -0.9 | 26.23 | 23.35 | 1.123 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 10.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.35 | TEVA | 2022-04-28 14:17 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.11 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.47 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:14 | GW012-Mannul-SP30-MD-01 | 1.043 | 0.043 | 450.6 | 6.1 | 214.1 | 0.02 | -12.9 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12.9 | -1.1 | 25.53 | 23.12 | 1.104 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 388 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 165.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.17 | TEVA | 2022-04-28 14:14 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.18 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.55 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:11 | GW012-Mannul-SP30-MD-01 | 1.053 | 0.052 | 499.2 | 6.8 | 272.1 | 0.283 | -11.5 | -0.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -11.5 | 0 | 26.43 | 23.95 | 1.104 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 370 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 3.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.6 | TEVA | 2022-04-28 14:11 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.3 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.32 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:08 | GW012-Mannul-SP30-MD-01 | 1.087 | 0.049 | 460.1 | 6.2 | 304.9 | 0.158 | -11.7 | 0.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -11.6 | 0.2 | 26.05 | 21.89 | 1.19 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 418 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 149.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.97 | TEVA | 2022-04-28 14:08 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.98 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.94 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:04 | GW012-Mannul-SP30-MD-01 | 1.065 | 0.057 | 435.9 | 5.9 | 223.9 | 0.196 | -10.3 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -10.4 | -1.2 | 24.85 | 21.57 | 1.152 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 378 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 140.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.07 | TEVA | 2022-04-28 14:04 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.61 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.44 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 14:02 | GW012-Mannul-SP30-MD-01 | 1.137 | 0.075 | 490.1 | 6.7 | 213 | 0.224 | -11.6 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -11.8 | -1.2 | 27.31 | 22.47 | 1.215 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 372 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 168.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.3 | TEVA | 2022-04-28 14:02 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.14 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.01 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:59 | GW012-Mannul-SP30-MD-01 | 1.055 | 0.054 | 477.6 | 6.5 | 153.4 | 0.131 | -9.4 | -0.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -9.6 | -0.2 | 25.99 | 23.24 | 1.118 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 424 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 116.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.06 | TEVA | 2022-04-28 13:59 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.74 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.05 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:56 | GW012-Mannul-SP30-MD-01 | 1.021 | 0.041 | 491.5 | 6.7 | 276.8 | 0.12 | -10.8 | 0 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -10.8 | 0.1 | 26.55 | 23.83 | 1.115 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 332 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 142.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.98 | TEVA | 2022-04-28 13:56 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.47 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.04 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:53 | GW012-Mannul-SP30-MD-01 | 1.162 | 0.062 | 454.6 | 6.2 | 261.4 | 0.203 | -11.4 | -0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -11.4 | -0.4 | 26.33 | 22.16 | 1.188 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 314 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 115.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.61 | TEVA | 2022-04-28 13:54 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.7 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:51 | GW012-Mannul-SP30-MD-01 | 1.135 | 0.055 | 487.4 | 6.6 | 268 | 0.189 | -11.5 | -0.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -11.5 | 0 | 27.18 | 22.62 | 1.201 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 424 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 30.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.24 | TEVA | 2022-04-28 13:51 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.12 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 79.14 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:48 | GW012-Mannul-SP30-MD-01 | 1.018 | 0.045 | 436 | 5.9 | 258.8 | 0.06 | -12.5 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -12.5 | -1.1 | 24.92 | 22.49 | 1.108 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 366 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 119.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.5 | TEVA | 2022-04-28 13:48 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.08 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.94 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:45 | GW012-Mannul-SP30-MD-01 | 1.127 | 0.062 | 443.9 | 6 | 177.8 | 0.068 | -14.5 | -2.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -14.5 | -2.5 | 25.87 | 21.79 | 1.187 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 294 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 98.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.95 | TEVA | 2022-04-28 13:45 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.13 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.21 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:42 | GW012-Mannul-SP30-MD-01 | 1.05 | 0.054 | 451 | 6.1 | 250.4 | 0.136 | -11.8 | -0.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -11.8 | -0.1 | 25.35 | 22.77 | 1.113 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 418 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 127.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.24 | TEVA | 2022-04-28 13:43 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.09 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.88 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 11 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:40 | GW012-Mannul-SP30-MD-01 | 1.095 | 0.027 | 428.4 | 5.8 | 231.7 | 0.066 | -10.7 | -0.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -10.7 | -0.3 | 24.95 | 21.98 | 1.135 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 416 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 93 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.14 | TEVA | 2022-04-28 13:40 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.4 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:37 | GW012-Mannul-SP30-MD-01 | 1.095 | 0.049 | 471.1 | 6.4 | 257.6 | 0.09 | -10.9 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -10.9 | -1.2 | 26.15 | 22.36 | 1.169 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 330 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 141 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.34 | TEVA | 2022-04-28 13:38 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.15 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.31 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:35 | GW012-Mannul-SP30-MD-01 | 1.06 | 0.036 | 407.1 | 5.5 | 315 | 0.108 | -11.8 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -11.7 | -0.8 | 24.09 | 21.59 | 1.116 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 418 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 118.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.14 | TEVA | 2022-04-28 13:35 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.83 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.2 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:32 | GW012-Mannul-SP30-MD-01 | 1.109 | 0.045 | 431.9 | 5.9 | 172.2 | 0.016 | -10.7 | -0.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -10.7 | -0.4 | 25.05 | 21.47 | 1.167 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 386 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 160.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.36 | TEVA | 2022-04-28 13:32 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.96 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.95 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:29 | GW012-Mannul-SP30-MD-01 | 1.157 | 0.038 | 483.9 | 6.6 | 271.3 | 0.143 | -10.6 | -0.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -10.6 | 0 | 26.89 | 22.44 | 1.198 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 164.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.35 | TEVA | 2022-04-28 13:30 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.78 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.3 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:27 | GW012-Mannul-SP30-MD-01 | 1.059 | 0.055 | 454.6 | 6.2 | 256.3 | 0.228 | -11.9 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -11.9 | -1.3 | 25.88 | 22.8 | 1.135 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 346 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 90 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.5 | TEVA | 2022-04-28 13:27 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.13 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.52 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 56 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:23 | GW012-Mannul-SP30-MD-01 | 1.08 | 0.041 | 426 | 5.8 | 170.3 | 0.105 | -12.1 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12.2 | -1.1 | 24.72 | 21.83 | 1.132 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 384 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 168.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.93 | TEVA | 2022-04-28 13:24 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.14 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.42 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:21 | GW012-Mannul-SP30-MD-01 | 1.062 | 0.034 | 417.7 | 5.7 | 283.8 | 0.023 | -12.5 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -12.5 | -0.9 | 24.46 | 21.96 | 1.113 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 414 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 44.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.46 | TEVA | 2022-04-28 13:21 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.1 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:18 | GW012-Mannul-SP30-MD-01 | 1.053 | 0.048 | 426.2 | 5.8 | 274.9 | 0.128 | -11.1 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -11.1 | -1.4 | 24.7 | 21.89 | 1.128 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 342 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 117.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.64 | TEVA | 2022-04-28 13:19 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.44 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.78 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 70 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:15 | GW012-Mannul-SP30-MD-01 | 1.036 | 0.053 | 456.2 | 6.2 | 357.9 | 0.108 | -10.1 | -0.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -10 | -0.2 | 24.89 | 22.8 | 1.092 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 384 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 161.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.25 | TEVA | 2022-04-28 13:16 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.4 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.24 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:13 | GW012-Mannul-SP30-MD-01 | 1.097 | 0.064 | 460.3 | 6.2 | 150.3 | 0.046 | -10.2 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -10.3 | -1.2 | 26.08 | 22.62 | 1.153 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 99.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.08 | TEVA | 2022-04-28 13:13 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.86 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.48 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:10 | GW012-Mannul-SP30-MD-01 | 1.106 | 0.048 | 381.9 | 5.2 | 258.4 | 0.179 | -10.8 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -10.8 | -0.8 | 23.75 | 19.71 | 1.205 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 396 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 115.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.91 | TEVA | 2022-04-28 13:11 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.72 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.89 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:08 | GW012-Mannul-SP30-MD-01 | 1.043 | 0.043 | 430.1 | 5.8 | 154.7 | 0.034 | -12.1 | -0.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12.1 | -0.1 | 24.62 | 22.14 | 1.112 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 400 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 123.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.61 | TEVA | 2022-04-28 13:08 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.64 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.88 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:05 | GW012-Mannul-SP30-MD-01 | 1.115 | 0.036 | 460.1 | 6.2 | 285.9 | 0.19 | -11.1 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -11 | -0.9 | 26.09 | 22.3 | 1.17 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 362 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 142 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.26 | TEVA | 2022-04-28 13:05 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.66 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.1 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 13:02 | GW012-Mannul-SP30-MD-01 | 1.076 | 0.042 | 449 | 6.1 | 114.4 | 0.028 | -12.6 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12.6 | -0.9 | 25.77 | 22.6 | 1.14 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 414 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 143.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.48 | TEVA | 2022-04-28 13:02 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.76 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.97 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 10 | 120 | cfpharmtech\mi1chen | 2022-04-28 12:59 | GW012-Mannul-SP30-MD-01 | 1.061 | 0.052 | 411.1 | 5.6 | 354.8 | 0.058 | -10.6 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -10.5 | -0.9 | 24.49 | 21.58 | 1.135 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 408 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 125.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.29 | TEVA | 2022-04-28 12:59 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.95 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.28 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 11:20 | GW012-Mannul-SP30-MD-01 | 1.063 | 0.043 | 434.3 | 5.9 | 293.7 | 0.078 | -11.8 | -0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -11.8 | -0.5 | 24.78 | 22.1 | 1.121 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 404 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 151.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.93 | TEVA | 2022-04-28 11:21 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.52 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.22 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 11:18 | GW012-Mannul-SP30-MD-01 | 1.058 | 0.067 | 508.1 | 6.9 | 230.9 | 0.315 | -11.9 | -0.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -12.1 | -0.2 | 26.77 | 23.81 | 1.124 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 316 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 7.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.86 | TEVA | 2022-04-28 11:18 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.44 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.4 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 11:15 | GW012-Mannul-SP30-MD-01 | 1.111 | 0.043 | 405.4 | 5.5 | 8.9 | 0.011 | -14 | -0.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -14 | -0.2 | 24.82 | 21.25 | 1.168 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 160.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.63 | TEVA | 2022-04-28 11:15 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.26 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.22 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 11:13 | GW012-Mannul-SP30-MD-01 | 1.042 | 0.062 | 406.4 | 5.5 | 187.8 | 0.063 | -11.3 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.4 | -0.8 | 23.71 | 21.52 | 1.101 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 382 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 83.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.93 | TEVA | 2022-04-28 11:13 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.01 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.84 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 11:10 | GW012-Mannul-SP30-MD-01 | 1.215 | 0.051 | 373.7 | 5.1 | 266.1 | 0.083 | -9.7 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -9.7 | -1.6 | 24.59 | 19.53 | 1.259 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 284 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 96.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.76 | TEVA | 2022-04-28 11:10 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 19.56 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.64 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 11:07 | GW012-Mannul-SP30-MD-01 | 1.072 | 0.051 | 424.9 | 5.8 | 257 | 0.152 | -10.7 | -0.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -10.8 | -0.2 | 25.03 | 21.58 | 1.16 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 386 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 136.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.8 | TEVA | 2022-04-28 11:08 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.19 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.55 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 11:05 | GW012-Mannul-SP30-MD-01 | 1.122 | 0.052 | 460.3 | 6.2 | 162 | 0.034 | -11.6 | -0.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.7 | -0.5 | 26.72 | 22.5 | 1.188 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 374 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 22.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.33 | TEVA | 2022-04-28 11:05 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.58 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.38 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 11:02 | GW012-Mannul-SP30-MD-01 | 1.039 | 0.068 | 410.8 | 5.6 | 187.6 | 0.058 | -13.3 | -0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -13.3 | -0.7 | 23.95 | 21.63 | 1.108 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 408 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 21.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.04 | TEVA | 2022-04-28 11:02 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.18 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.05 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 10:58 | GW012-Mannul-SP30-MD-01 | 1.08 | 0.062 | 415.1 | 5.6 | 325.9 | 0.053 | -12.1 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -12.1 | -0.9 | 24.41 | 21.58 | 1.131 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 71.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.58 | TEVA | 2022-04-28 10:58 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.83 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.12 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 10:56 | GW012-Mannul-SP30-MD-01 | 1.063 | 0.033 | 391.3 | 5.3 | 341.2 | 0.116 | -11.9 | -2.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.8 | -2.2 | 24 | 21.19 | 1.133 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 340 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 155.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.7 | TEVA | 2022-04-28 10:56 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.35 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.09 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 10:53 | GW012-Mannul-SP30-MD-01 | 1.063 | 0.086 | 431.6 | 5.9 | 230.4 | 0.132 | -11 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -11.1 | -1.1 | 24.88 | 22.34 | 1.114 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 358 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 8.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.94 | TEVA | 2022-04-28 10:53 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.52 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.48 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 10:50 | GW012-Mannul-SP30-MD-01 | 1.065 | 0.032 | 418.5 | 5.7 | 184.6 | 0.164 | -11.5 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -11.7 | -1.6 | 24.46 | 21.48 | 1.138 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 372 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 153.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.52 | TEVA | 2022-04-28 10:51 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.07 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 50 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 10:47 | GW012-Mannul-SP30-MD-01 | 1.1 | 0.042 | 488.1 | 6.6 | 189 | 0.192 | -14.9 | -0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -15.1 | -0.6 | 26.99 | 23 | 1.174 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 322 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 173.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.86 | TEVA | 2022-04-28 10:47 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.49 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 79.57 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 70 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 10:44 | GW012-Mannul-SP30-MD-01 | 1.12 | 0.074 | 409.1 | 5.6 | 46.6 | 0.126 | -12.1 | -0.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -12.1 | -0.6 | 24.42 | 20.78 | 1.175 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 386 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 122.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.87 | TEVA | 2022-04-28 10:44 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.31 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.87 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 9 | 120 | cfpharmtech\mi1chen | 2022-04-28 10:40 | GW012-Mannul-SP30-MD-01 | 1.103 | 0.065 | 442.8 | 6 | 76.7 | 0.126 | -11.2 | -0.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -11.2 | -0.6 | 25.7 | 21.34 | 1.205 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 410 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 45.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.69 | TEVA | 2022-04-28 10:40 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.39 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.67 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 10:01 | GW012-Mannul-SP30-MD-01 | 1.109 | 0.08 | 484.9 | 6.6 | 260.2 | 0.101 | -12.6 | 0 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.6 | 0.1 | 27 | 22.6 | 1.194 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 42.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.89 | TEVA | 2022-04-28 10:01 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.35 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.34 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 82 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:57 | GW012-Mannul-SP30-MD-01 | 1.059 | 0.044 | 390.9 | 5.3 | 216.4 | 0.022 | -12.1 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.1 | -1.1 | 23.64 | 21.11 | 1.12 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 412 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 37.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.66 | TEVA | 2022-04-28 09:58 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.39 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.88 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:54 | GW012-Mannul-SP30-MD-01 | 1.142 | 0.056 | 447.7 | 6.1 | 132.6 | 0.039 | -11.2 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -11.2 | -1.3 | 26.51 | 22.04 | 1.203 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 382 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 177.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.22 | TEVA | 2022-04-28 09:54 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.09 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.06 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:52 | GW012-Mannul-SP30-MD-01 | 1.141 | 0.06 | 436.2 | 5.9 | 226.3 | 0.222 | -10.6 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -10.7 | -1 | 26.23 | 21.44 | 1.223 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 344 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 116.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.85 | TEVA | 2022-04-28 09:52 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.79 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.29 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 54 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:49 | GW012-Mannul-SP30-MD-01 | 1.048 | 0.068 | 450.5 | 6.1 | 136.5 | 0.027 | -12.7 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -12.7 | -1 | 25.66 | 22.49 | 1.141 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 386 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 79.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.24 | TEVA | 2022-04-28 09:49 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.14 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.34 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:47 | GW012-Mannul-SP30-MD-01 | 1.057 | 0.034 | 389.3 | 5.3 | 26 | 0.038 | -11.4 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -11.3 | -1.6 | 23.92 | 21.1 | 1.134 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 330 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 153.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.58 | TEVA | 2022-04-28 09:47 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.36 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.74 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:44 | GW012-Mannul-SP30-MD-01 | 1.095 | 0.055 | 435.3 | 5.9 | 274.4 | 0.2 | -9.2 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.4 | MDx |  |  |  |  | -9.2 | -1 | 25.79 | 21.86 | 1.18 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 328 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 104.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.34 | TEVA | 2022-04-28 09:44 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.22 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.65 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 70 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:40 | GW012-Mannul-SP30-MD-01 | 1.096 | 0.069 | 417.8 | 5.7 | 216.6 | 0.121 | -9.5 | -2.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.3 | MDx |  |  |  |  | -9.6 | -2.2 | 24.91 | 21 | 1.186 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 348 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 13.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.81 | TEVA | 2022-04-28 09:40 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.73 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.95 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 54 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:37 | GW012-Mannul-SP30-MD-01 | 1.133 | 0.061 | 405.4 | 5.5 | 201.7 | 0.284 | -14.3 | -2.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.1 | MDx |  |  |  |  | -14.6 | -2.5 | 24.51 | 19.96 | 1.228 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 292 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 140.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.92 | TEVA | 2022-04-28 09:37 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.11 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.73 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:34 | GW012-Mannul-SP30-MD-01 | 1.07 | 0.045 | 476.7 | 6.5 | 205.6 | 0.144 | -11 | -0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36 | MDx |  |  |  |  | -11.1 | -0.6 | 26.09 | 22.96 | 1.137 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 360 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 22.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.18 | TEVA | 2022-04-28 09:34 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.52 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 79.09 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 98 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:31 | GW012-Mannul-SP30-MD-01 | 1.045 | 0.063 | 390 | 5.3 | 23.9 | 0.031 | -10.1 | -0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36 | MDx |  |  |  |  | -10.1 | -0.7 | 23.33 | 21.17 | 1.102 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 422 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 171.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.52 | TEVA | 2022-04-28 09:31 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.55 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.83 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 82 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:28 | GW012-Mannul-SP30-MD-01 | 1.038 | 0.047 | 399.7 | 5.4 | 294.1 | 0.082 | -10.1 | -0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 35.7 | MDx |  |  |  |  | -10 | -0.7 | 23.67 | 21.71 | 1.09 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 412 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 153 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.69 | TEVA | 2022-04-28 09:28 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.86 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.59 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:25 | GW012-Mannul-SP30-MD-01 | 1.047 | 0.036 | 401.6 | 5.5 | 196.2 | 0.099 | -10.2 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 35.4 | MDx |  |  |  |  | -10.3 | -1.2 | 23.81 | 21.44 | 1.111 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 324 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 146.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.83 | TEVA | 2022-04-28 09:25 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.8 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.04 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:22 | GW012-Mannul-SP30-MD-01 | 1.039 | 0.06 | 379.6 | 5.2 | 107.1 | 0.042 | -9.9 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 35.3 | MDx |  |  |  |  | -9.9 | -1.2 | 23.25 | 21 | 1.107 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 386 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 112.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.11 | TEVA | 2022-04-28 09:22 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.28 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.66 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 82 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 8 | 120 | cfpharmtech\mi1chen | 2022-04-28 09:19 | GW012-Mannul-SP30-MD-01 | 1.103 | 0.04 | 403.4 | 5.5 | 282.3 | 0.126 | -9.8 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 35.1 | MDx |  |  |  |  | -9.8 | -0.9 | 24.09 | 21 | 1.148 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 125.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.5 | TEVA | 2022-04-28 09:19 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.31 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.91 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:57 | GW012-Mannul-SP30-MD-01 | 1.084 | 0.038 | 374.7 | 5 | 287.9 | 0.195 | -13 | 0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -12.9 | 0.9 | 23.74 | 20.7 | 1.147 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 52.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.45 | TEVA | 2022-04-27 16:57 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.71 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.16 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 102 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:51 | GW012-Mannul-SP30-MD-01 | 1.031 | 0.048 | 403.7 | 5.4 | 275 | 0.133 | -12.7 | 1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -12.7 | 1.7 | 23.78 | 21.62 | 1.1 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 400 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 77.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.73 | TEVA | 2022-04-27 16:53 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.05 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.1 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 80 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:48 | GW012-Mannul-SP30-MD-01 | [8 empty] | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx | [9 empty] | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 498 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 |  | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 |  | TEVA | 2022-04-27 16:53 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 |  | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 |  | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 4 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:45 | GW012-Mannul-SP30-MD-01 | 1.029 | 0.053 | 353.4 | 4.7 | 87.4 | 0.073 | -11.8 | 1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.8 | 1 | 22.23 | 20.29 | 1.096 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 402 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 64.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 21.2 | TEVA | 2022-04-27 16:45 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.59 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 67.05 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:42 | GW012-Mannul-SP30-MD-01 | 1.095 | 0.047 | 356.3 | 4.8 | 321.1 | 0.079 | -12 | 1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.9 | 1.4 | 22.99 | 19.97 | 1.151 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 382 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 165.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 21.98 | TEVA | 2022-04-27 16:43 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.07 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 67.61 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:39 | GW012-Mannul-SP30-MD-01 | 1.103 | 0.057 | 390 | 5.2 | 181.5 | 0.14 | -12.3 | 1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -12.5 | 1.1 | 23.79 | 20.7 | 1.15 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 316 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 102.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.1 | TEVA | 2022-04-27 16:39 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.94 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.81 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 74 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:34 | GW012-Mannul-SP30-MD-01 | 1.106 | 0.068 | 380 | 5.1 | 229.1 | 0.302 | -13.6 | -0.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -13.8 | 0 | 23.63 | 20.43 | 1.156 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 382 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 17.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.8 | TEVA | 2022-04-27 16:34 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.62 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.33 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 58 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:31 | GW012-Mannul-SP30-MD-01 | 1.096 | 0.057 | 374.3 | 5 | 235.1 | 0.121 | -13.7 | 0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -13.8 | 1 | 23.31 | 20.16 | 1.156 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 352 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 164.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.6 | TEVA | 2022-04-27 16:31 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.62 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.69 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:16 | GW012-Mannul-SP30-MD-01 | 1.068 | 0.036 | 373 | 5 | 295.7 | 0.121 | -12.7 | 0.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -12.7 | 0.5 | 22.99 | 20.12 | 1.143 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 366 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 144.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.22 | TEVA | 2022-04-27 16:16 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.8 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 68.36 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:13 | GW012-Mannul-SP30-MD-01 | 1.042 | 0.052 | 367.2 | 4.9 | 180.7 | 0.08 | -12.3 | 0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -12.4 | 0.7 | 22.9 | 20.71 | 1.106 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 404 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 41.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 21.75 | TEVA | 2022-04-27 16:13 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.87 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.15 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:10 | GW012-Mannul-SP30-MD-01 | 1.168 | 0.079 | 423.1 | 5.6 | 213.2 | 0.165 | -13.2 | 0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -13.4 | 0.7 | 25.46 | 20.96 | 1.215 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 360 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 1.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.78 | TEVA | 2022-04-27 16:10 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.21 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.37 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 56 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:07 | GW012-Mannul-SP30-MD-01 | 1.052 | 0.034 | 356 | 4.8 | 289 | 0.134 | -13.3 | 7.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -13.2 | 8.1 | 22.6 | 20.32 | 1.112 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 410 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 66.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 21.53 | TEVA | 2022-04-27 16:07 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.46 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 67.86 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:03 | GW012-Mannul-SP30-MD-01 | 1.156 | 0.062 | 474.8 | 6.3 | 199.9 | 0.097 | -12.2 | 1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -12.3 | 1.5 | 27.13 | 22.23 | 1.22 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 316 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 143.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.13 | TEVA | 2022-04-27 16:04 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.61 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 79.04 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 74 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 16:00 | GW012-Mannul-SP30-MD-01 | 1.057 | 0.054 | 396.5 | 5.3 | 235.7 | 0.094 | -12.2 | 1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -12.2 | 1.7 | 23.94 | 20.86 | 1.148 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 414 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 53.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.8 | TEVA | 2022-04-27 16:00 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.57 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.45 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 7 | 120 | cfpharmtech\mi1chen | 2022-04-27 15:57 | GW012-Mannul-SP30-MD-01 | 1.057 | 0.057 | 317.7 | 4.2 | 350.8 | 0.126 | -10.7 | 0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -10.5 | 0.9 | 21.22 | 18.86 | 1.125 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 412 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 64.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 20.37 | TEVA | 2022-04-27 15:57 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 19.27 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 64 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 98 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 15:19 | GW012-Mannul-SP30-MD-01 | 1.045 | 0.065 | 430.6 | 5.7 | 218.8 | 0.072 | -12.4 | 0.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -12.5 | 0.3 | 24.4 | 22 | 1.109 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 370 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 146.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.67 | TEVA | 2022-04-27 15:19 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.65 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.52 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 76 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 15:16 | GW012-Mannul-SP30-MD-01 | 1.066 | 0.041 | 429.7 | 5.7 | 63.1 | 0.083 | -14.3 | 1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -14.2 | 0.9 | 25.02 | 21.92 | 1.142 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 436 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 169.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.85 | TEVA | 2022-04-27 15:17 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.37 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.36 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 15:14 | GW012-Mannul-SP30-MD-01 | 1.044 | 0.039 | 410.8 | 5.5 | 218.7 | 0.066 | -13.5 | 0.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.1 | MDx |  |  |  |  | -13.5 | 0.5 | 24.24 | 21.75 | 1.114 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 408 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 161.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.06 | TEVA | 2022-04-27 15:14 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.08 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.91 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 15:11 | GW012-Mannul-SP30-MD-01 | 1.119 | 0.045 | 384.2 | 5.1 | 186 | 0.111 | -13.6 | 0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.2 | MDx |  |  |  |  | -13.7 | 0.6 | 23.67 | 20.17 | 1.173 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 384 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 173.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.05 | TEVA | 2022-04-27 15:11 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.6 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.16 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 98 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 15:07 | GW012-Mannul-SP30-MD-01 | 1.104 | 0.067 | 437 | 5.8 | 231 | 0.3 | -13.7 | 1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.3 | MDx |  |  |  |  | -13.9 | 1.8 | 25.8 | 21.75 | 1.186 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 306 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 69.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.53 | TEVA | 2022-04-27 15:07 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.21 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.94 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 15:04 | GW012-Mannul-SP30-MD-01 | 1.073 | 0.057 | 345.4 | 4.6 | 44.3 | 0.055 | -12 | 0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.3 | MDx |  |  |  |  | -12 | 0.7 | 22.33 | 19.88 | 1.123 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 428 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 16.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 21.4 | TEVA | 2022-04-27 15:04 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 19.94 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 66.39 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 15:00 | GW012-Mannul-SP30-MD-01 | 1.03 | 0.049 | 415.7 | 5.6 | 196.7 | 0.088 | -11.9 | 0.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.3 | MDx |  |  |  |  | -12 | 0.4 | 24.07 | 21.86 | 1.101 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 388 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 54.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.08 | TEVA | 2022-04-27 15:00 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.41 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.01 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 102 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:57 | GW012-Mannul-SP30-MD-01 | 1.225 | 0.069 | 415.2 | 5.5 | 231.2 | 0.137 | -12.5 | 0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.3 | MDx |  |  |  |  | -12.6 | 0.9 | 25.96 | 20.5 | 1.266 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 342 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 176.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.15 | TEVA | 2022-04-27 14:57 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.54 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.72 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:54 | GW012-Mannul-SP30-MD-01 | 1.068 | 0.042 | 380.9 | 5.1 | 236.9 | 0.242 | -11.2 | 0 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.6 | MDx |  |  |  |  | -11.3 | 0.3 | 23.59 | 20.83 | 1.133 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 276 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 106.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.45 | TEVA | 2022-04-27 14:54 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.02 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.48 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:50 | GW012-Mannul-SP30-MD-01 | 1.128 | 0.075 | 388.5 | 5.2 | 207.1 | 0.109 | -12.8 | 0.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.6 | MDx |  |  |  |  | -12.9 | 0.5 | 24.88 | 20.42 | 1.218 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 402 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 36.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.3 | TEVA | 2022-04-27 14:50 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.65 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.55 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:47 | GW012-Mannul-SP30-MD-01 | 1.025 | 0.079 | 478.9 | 6.4 | 226.5 | 0.38 | -13.7 | 1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.7 | MDx |  |  |  |  | -13.9 | 1.9 | 26.67 | 23.29 | 1.145 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 306 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 125.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.76 | TEVA | 2022-04-27 14:47 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 24.15 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 79.41 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:44 | GW012-Mannul-SP30-MD-01 | 1.042 | 0.043 | 402 | 5.4 | 101.3 | 0.032 | -13.2 | 0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.7 | MDx |  |  |  |  | -13.2 | 0.6 | 23.74 | 21.35 | 1.112 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 420 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 105.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.77 | TEVA | 2022-04-27 14:44 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.87 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.16 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:41 | GW012-Mannul-SP30-MD-01 | 1.133 | 0.062 | 399.4 | 5.3 | 198.9 | 0.245 | -13.7 | 0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.9 | MDx |  |  |  |  | -14 | 0.7 | 24.42 | 20.35 | 1.2 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 157.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.74 | TEVA | 2022-04-27 14:42 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.95 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.01 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:38 | GW012-Mannul-SP30-MD-01 | 1.045 | 0.054 | 432.1 | 5.8 | 247 | 0.116 | -13.2 | 1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.9 | MDx |  |  |  |  | -13.3 | 1.7 | 24.48 | 22.41 | 1.093 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 396 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 94.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.68 | TEVA | 2022-04-27 14:38 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.66 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.85 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 6 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:35 | GW012-Mannul-SP30-MD-01 | 1.098 | 0.045 | 386.4 | 5.2 | 49.9 | 0.02 | -11.8 | 0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.7 | MDx |  |  |  |  | -11.8 | 0.9 | 23.54 | 20.74 | 1.135 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 414 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 73.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.92 | TEVA | 2022-04-27 14:35 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.88 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.78 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:25 | GW012-Mannul-SP30-MD-01 | 1.085 | 0.04 | 419.6 | 5.6 | 247.2 | 0.116 | -10.2 | -0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.2 | MDx |  |  |  |  | -10.3 | -0.5 | 24.71 | 21.29 | 1.161 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 290 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 94.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.76 | TEVA | 2022-04-27 14:26 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.9 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.3 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 76 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:22 | GW012-Mannul-SP30-MD-01 | 1.155 | 0.047 | 351.3 | 4.7 | 155.8 | 0.026 | -11.5 | 0.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.3 | MDx |  |  |  |  | -11.5 | 0.1 | 23.13 | 19.11 | 1.21 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 87.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.41 | TEVA | 2022-04-27 14:22 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 19.4 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 67.22 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:19 | GW012-Mannul-SP30-MD-01 | 1.137 | 0.044 | 409.8 | 5.5 | 247.9 | 0.227 | -14.2 | 0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.3 | MDx |  |  |  |  | -14.3 | 0.9 | 25.24 | 20.7 | 1.22 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 408 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 155.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.05 | TEVA | 2022-04-27 14:20 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.15 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.99 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:16 | GW012-Mannul-SP30-MD-01 | 1.133 | 0.048 | 441.5 | 5.9 | 209.8 | 0.385 | -12.7 | 0.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.6 | MDx |  |  |  |  | -13.1 | 0.5 | 26.04 | 21.85 | 1.192 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 352 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 152.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.9 | TEVA | 2022-04-27 14:16 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.98 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.12 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:13 | GW012-Mannul-SP30-MD-01 | 1.145 | 0.068 | 397.6 | 5.3 | 261 | 0.301 | -12 | 0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.8 | MDx |  |  |  |  | -12.1 | 0.9 | 25.23 | 20.95 | 1.204 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 332 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 96.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.81 | TEVA | 2022-04-27 14:13 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.78 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.23 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:10 | GW012-Mannul-SP30-MD-01 | 1.055 | 0.035 | 386.4 | 5.2 | 288.8 | 0.062 | -12.7 | 0.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.9 | MDx |  |  |  |  | -12.6 | 0.6 | 23.43 | 20.86 | 1.123 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 422 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 55 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.47 | TEVA | 2022-04-27 14:10 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.31 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.85 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 80 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:07 | GW012-Mannul-SP30-MD-01 | 1.016 | 0.054 | 461.2 | 6.2 | 251.2 | 0.277 | -12.7 | 1.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.9 | MDx |  |  |  |  | -12.8 | 1.7 | 25.3 | 22.97 | 1.102 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 330 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 125.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.15 | TEVA | 2022-04-27 14:07 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.78 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.33 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 14:03 | GW012-Mannul-SP30-MD-01 | 1.086 | 0.044 | 413 | 5.5 | 202.2 | 0.103 | -12.8 | 1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 40.2 | MDx |  |  |  |  | -12.9 | 1.2 | 24.34 | 21.41 | 1.137 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 372 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 69.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.57 | TEVA | 2022-04-27 14:03 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.71 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.53 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 13:59 | GW012-Mannul-SP30-MD-01 | 1.035 | 0.045 | 404.8 | 5.4 | 256.4 | 0.092 | -13.5 | 0.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 40.3 | MDx |  |  |  |  | -13.5 | 0.4 | 23.62 | 21.64 | 1.092 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 412 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 104.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.78 | TEVA | 2022-04-27 13:59 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.01 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.2 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 13:56 | GW012-Mannul-SP30-MD-01 | 1.026 | 0.051 | 432.4 | 5.8 | 207.8 | 0.072 | -12.9 | 0.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 40.5 | MDx |  |  |  |  | -12.9 | 0.4 | 24.53 | 22.22 | 1.104 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 378 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 14.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.49 | TEVA | 2022-04-27 13:56 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.89 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.15 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 100 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 13:53 | GW012-Mannul-SP30-MD-01 | 1.1 | 0.047 | 406.8 | 5.4 | 136.7 | 0.049 | -12.5 | 0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 40.6 | MDx |  |  |  |  | -12.5 | 0.7 | 24.53 | 21.18 | 1.158 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 400 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 20.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.57 | TEVA | 2022-04-27 13:53 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.43 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.57 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 80 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 13:50 | GW012-Mannul-SP30-MD-01 | 1.239 | 0.072 | 403.7 | 5.4 | 224.7 | 0.206 | -12 | 2.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 40.8 | MDx |  |  |  |  | -12.2 | 2.4 | 25.56 | 19.59 | 1.305 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 338 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 115.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.89 | TEVA | 2022-04-27 13:51 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.1 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.1 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 13:47 | GW012-Mannul-SP30-MD-01 | 1.066 | 0.051 | 404.5 | 5.4 | 261.1 | 0.125 | -12.9 | 0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 41.2 | MDx |  |  |  |  | -12.9 | 0.7 | 24.02 | 20.62 | 1.165 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 356 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 95.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.11 | TEVA | 2022-04-27 13:48 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.68 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.15 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 13:44 | GW012-Mannul-SP30-MD-01 | 1.061 | 0.054 | 414.7 | 5.5 | 335.3 | 0.081 | -11.7 | 1.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 41.5 | MDx |  |  |  |  | -11.6 | 1.4 | 24.36 | 21.77 | 1.119 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 398 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 106.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.39 | TEVA | 2022-04-27 13:45 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.02 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 5 | 120 | cfpharmtech\mi1chen | 2022-04-27 13:41 | GW012-Mannul-SP30-MD-01 | 1.094 | 0.039 | 404.4 | 5.4 | 286.7 | 0.032 | -13.4 | 1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 41.8 | MDx |  |  |  |  | -13.4 | 1.1 | 24.04 | 21.41 | 1.123 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 89.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.42 | TEVA | 2022-04-27 13:41 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.41 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.78 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:59 | GW012-Mannul-SP30-MD-01 | 1.164 | 0.061 | 447.6 | 6 | 231.3 | 0.305 | -11.5 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -11.7 | -1.4 | 26.06 | 21.06 | 1.237 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 360 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 142.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.48 | TEVA | 2022-04-26 14:59 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.9 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.31 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:56 | GW012-Mannul-SP30-MD-01 | 1.145 | 0.047 | 378 | 5.1 | 233.2 | 0.075 | -11.4 | -3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -11.4 | -2.9 | 23.64 | 19.74 | 1.198 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 310 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 115.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.15 | TEVA | 2022-04-26 14:56 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.22 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.21 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:53 | GW012-Mannul-SP30-MD-01 | 1.137 | 0.052 | 495.8 | 6.7 | 219.2 | 0.255 | -10.3 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -10.5 | -0.6 | 27.08 | 22.8 | 1.188 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 400 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 4.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.46 | TEVA | 2022-04-26 14:53 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.28 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.3 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:50 | GW012-Mannul-SP30-MD-01 | 1.039 | 0.042 | 468.5 | 6.3 | 195.8 | 0.011 | -11.4 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -11.4 | -1.2 | 25.88 | 23.28 | 1.111 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 402 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 120.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.61 | TEVA | 2022-04-26 14:50 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.68 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.3 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:47 | GW012-Mannul-SP30-MD-01 | 1.156 | 0.063 | 424.5 | 5.7 | 289 | 0.198 | -10.4 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.8 | MDx |  |  |  |  | -10.3 | -1.6 | 25.46 | 21.07 | 1.208 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 322 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 113.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.72 | TEVA | 2022-04-26 14:47 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.39 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.09 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:44 | GW012-Mannul-SP30-MD-01 | 1.108 | 0.057 | 452.9 | 6.1 | 159.1 | 0.235 | -11.8 | -2.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -12 | -2.2 | 25.75 | 21.59 | 1.193 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 402 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 15.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.94 | TEVA | 2022-04-26 14:44 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.51 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:40 | GW012-Mannul-SP30-MD-01 | 1.054 | 0.054 | 473.8 | 6.4 | 223.5 | 0.209 | -10.3 | -0.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -10.5 | -0.3 | 25.79 | 22.62 | 1.14 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 320 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 113.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.97 | TEVA | 2022-04-26 14:40 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.69 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.25 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:37 | GW012-Mannul-SP30-MD-01 | 1.073 | 0.056 | 472.8 | 6.4 | 223.7 | 0.125 | -11.2 | -0.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -11.3 | -0.5 | 26.06 | 23.21 | 1.123 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 390 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 8.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.1 | TEVA | 2022-04-26 14:37 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.4 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.21 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:34 | GW012-Mannul-SP30-MD-01 | 1.117 | 0.043 | 495.5 | 6.7 | 249.5 | 0.153 | -12 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -12 | -1.2 | 27.22 | 23.17 | 1.175 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 322 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 121.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.22 | TEVA | 2022-04-26 14:34 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.47 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 80.83 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 72 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:31 | GW012-Mannul-SP30-MD-01 | 1.004 | 0.056 | 452.6 | 6.1 | 45.8 | 0.138 | -11 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -10.9 | -1.4 | 25.43 | 22.87 | 1.112 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 396 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 162.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.73 | TEVA | 2022-04-26 14:31 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.63 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.9 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:28 | GW012-Mannul-SP30-MD-01 | 1.07 | 0.069 | 430.5 | 5.8 | 214.5 | 0.098 | -10.2 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -10.3 | -1.2 | 24.7 | 21.99 | 1.123 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 370 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 14.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.95 | TEVA | 2022-04-26 14:28 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.37 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.56 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 100 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:24 | GW012-Mannul-SP30-MD-01 | 1.111 | 0.061 | 433.1 | 5.8 | 118.1 | 0.116 | -10.7 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -10.8 | -1.2 | 25 | 21.68 | 1.153 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 388 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 17.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.43 | TEVA | 2022-04-26 14:25 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.99 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.14 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:22 | GW012-Mannul-SP30-MD-01 | 1.062 | 0.063 | 412.7 | 5.6 | 335.4 | 0.154 | -11 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -10.9 | -0.9 | 24.26 | 21.88 | 1.109 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 390 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 150 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.34 | TEVA | 2022-04-26 14:22 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.98 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.4 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:18 | GW012-Mannul-SP30-MD-01 | 1.126 | 0.049 | 451.1 | 6.1 | 242.8 | 0.269 | -10.8 | -1.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -10.9 | -1.6 | 26.36 | 21.66 | 1.217 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 362 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 133.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.14 | TEVA | 2022-04-26 14:19 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.32 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.12 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 4 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:15 | GW012-Mannul-SP30-MD-01 | 1.026 | 0.036 | 386.6 | 5.2 | 191.4 | 0.176 | -12.3 | -1.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.4 | -1.7 | 23.42 | 20.94 | 1.118 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 388 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 146.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.18 | TEVA | 2022-04-26 14:15 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.61 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.08 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 104 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:12 | GW012-Mannul-SP30-MD-01 | 1.061 | 0.033 | 450.3 | 6.1 | 197.8 | 0.063 | -12.5 | -0.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.5 | -0.7 | 25.42 | 22.87 | 1.112 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 424 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 173.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.39 | TEVA | 2022-04-26 14:12 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.98 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.26 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:09 | GW012-Mannul-SP30-MD-01 | 1.041 | 0.065 | 421 | 5.7 | 305.7 | 0.212 | -11.3 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -11.2 | -0.9 | 24.37 | 21.96 | 1.11 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 488 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 116.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.34 | TEVA | 2022-04-26 14:09 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.42 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.26 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:04 | GW012-Mannul-SP30-MD-01 | 1.113 | 0.052 | 476.7 | 6.4 | 266.6 | 0.253 | -11 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -11 | -1.5 | 26.64 | 22.44 | 1.187 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 336 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 139.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.69 | TEVA | 2022-04-26 14:05 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.09 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.61 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 14:01 | GW012-Mannul-SP30-MD-01 | 1.174 | 0.063 | 453.8 | 6.1 | 206.2 | 0.109 | -12.5 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.6 | -1.3 | 26.62 | 21.38 | 1.245 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 370 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 168.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.75 | TEVA | 2022-04-26 14:01 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.93 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.94 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 13:57 | GW012-Mannul-SP30-MD-01 | 1.051 | 0.056 | 410 | 5.5 | 304.6 | 0.102 | -9.5 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -9.5 | -1 | 24.36 | 21.23 | 1.147 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 382 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 87.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.09 | TEVA | 2022-04-26 13:57 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.96 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.31 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 13:54 | GW012-Mannul-SP30-MD-01 | 1.034 | 0.062 | 435.6 | 5.9 | 200.8 | 0.071 | -10.1 | -2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -10.2 | -2 | 24.74 | 22.13 | 1.118 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 390 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 131.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.64 | TEVA | 2022-04-26 13:54 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.86 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.49 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 13:50 | GW012-Mannul-SP30-MD-01 | 1.196 | 0.057 | 502.6 | 6.8 | 206.8 | 0.299 | -11 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -11.3 | -1.4 | 28.43 | 22.5 | 1.263 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 392 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 25.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 27.32 | TEVA | 2022-04-26 13:50 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.85 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 81.59 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 102 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 13:43 | GW012-Mannul-SP30-MD-01 | 1.818 | 0.066 | 707.1 | 9.5 | 270.9 | 0.228 | 2.1 | -10.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | 2.1 | -10.4 | 41.57 | 21.92 | 1.896 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 380 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 1.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 40.08 | TEVA | 2022-04-26 13:43 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 102.67 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 13:39 | GW012-Mannul-SP30-MD-01 | 1.863 | 0.043 | 675.9 | 9.1 | 292.1 | 0.317 | 4.2 | -10.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | 4.3 | -10.6 | 40.86 | 20.97 | 1.949 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 360 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 174 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 39.64 | TEVA | 2022-04-26 13:39 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.28 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 98.35 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 54 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 13:23 | GW012-Mannul-SP30-MD-01 | 1.053 | 0.045 | 406.3 | 5.4 | 357.1 | 0.136 | -12.2 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -12 | -1.7 | 23.99 | 21.62 | 1.11 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 368 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 138.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.03 | TEVA | 2022-04-26 13:23 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.86 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.15 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 98 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 13:19 | GW012-Mannul-SP30-MD-01 | 1.149 | 0.057 | 424.2 | 5.7 | 231 | 0.256 | -12.8 | -1.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37 | MDx |  |  |  |  | -12.9 | -1.2 | 25.47 | 20.6 | 1.236 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 344 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 112.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.65 | TEVA | 2022-04-26 13:19 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.45 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.35 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 13:14 | GW012-Mannul-SP30-MD-01 | 1.042 | 0.067 | 446.8 | 6 | 168.6 | 0.038 | -12.3 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -12.3 | -1.7 | 25.16 | 22.92 | 1.098 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 304 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 129.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.06 | TEVA | 2022-04-26 13:14 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.1 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.6 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 74 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 13:09 | GW012-Mannul-SP30-MD-01 | 1.206 | 0.046 | 397.2 | 5.3 | 283.9 | 0.158 | -10.9 | -1.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.7 | MDx |  |  |  |  | -10.8 | -1.1 | 25.28 | 20.08 | 1.259 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 51.4 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.36 | TEVA | 2022-04-26 13:09 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.2 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.64 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 13:06 | GW012-Mannul-SP30-MD-01 | 1.108 | 0.033 | 507.5 | 6.8 | 221.7 | 0.354 | -10.8 | -0.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 36.6 | MDx |  |  |  |  | -11.1 | -0.1 | 27.16 | 23.84 | 1.14 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 326 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 146.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 26.44 | TEVA | 2022-04-26 13:06 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.86 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 81.13 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 70 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 3 | 120 | cfpharmtech\mi1chen | 2022-04-26 11:03 | GW012-Mannul-SP30-MD-01 | 1.105 | 0.077 | 387.5 | 5.2 | 216.3 | 0.142 | -11.2 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.2 | MDx |  |  |  |  | -11.3 | -1.3 | 23.76 | 20.57 | 1.155 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 402 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 157.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.07 | TEVA | 2022-04-26 11:03 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.88 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 70.92 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:58 | GW012-Mannul-SP30-MD-01 | 1.056 | 0.049 | 440 | 5.9 | 341.9 | 0.101 | -11.2 | -0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -11.1 | -0.6 | 25.09 | 22.12 | 1.135 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 380 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 147.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.03 | TEVA | 2022-04-26 10:58 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.76 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.83 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:55 | GW012-Mannul-SP30-MD-01 | 1.053 | 0.046 | 409.3 | 5.5 | 229.6 | 0.128 | -11.9 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -12 | -0.9 | 24.03 | 21.5 | 1.117 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 410 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 119.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.14 | TEVA | 2022-04-26 10:55 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.98 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.92 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 96 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:52 | GW012-Mannul-SP30-MD-01 | 1.081 | 0.059 | 463.2 | 6.2 | 220.1 | 0.101 | -11.7 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.4 | MDx |  |  |  |  | -11.8 | -0.9 | 25.42 | 22.52 | 1.129 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 398 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 3.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.92 | TEVA | 2022-04-26 10:52 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.05 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.01 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:49 | GW012-Mannul-SP30-MD-01 | 1.06 | 0.062 | 376.2 | 5 | 271 | 0.065 | -10.8 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -10.8 | -1.3 | 23.4 | 20.69 | 1.131 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 378 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 124.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 22.22 | TEVA | 2022-04-26 10:49 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 20.96 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 69.11 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:46 | GW012-Mannul-SP30-MD-01 | 1.132 | 0.063 | 443.5 | 5.9 | 241.3 | 0.187 | -11.7 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.6 | MDx |  |  |  |  | -11.8 | -0.8 | 26.29 | 22.42 | 1.173 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 428 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 170.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.03 | TEVA | 2022-04-26 10:46 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.1 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.36 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:43 | GW012-Mannul-SP30-MD-01 | 1.095 | 0.055 | 413.7 | 5.5 | 201.6 | 0.189 | -10.4 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -10.6 | -1.4 | 24.83 | 21.15 | 1.174 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 338 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.7 | TEVA | 2022-04-26 10:43 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.64 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.57 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:39 | GW012-Mannul-SP30-MD-01 | 1.037 | 0.055 | 450.9 | 6 | 199.9 | 0.17 | -10.6 | -0.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.7 | MDx |  |  |  |  | -10.7 | -0.3 | 25.09 | 22.53 | 1.114 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 342 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 145.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.13 | TEVA | 2022-04-26 10:39 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.26 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.02 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:36 | GW012-Mannul-SP30-MD-01 | 1.092 | 0.055 | 435 | 5.8 | 310.6 | 0.1 | -12.2 | -0.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -12.1 | -0.5 | 25.54 | 22.08 | 1.157 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 420 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 2.9 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.28 | TEVA | 2022-04-26 10:36 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.24 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.6 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:32 | GW012-Mannul-SP30-MD-01 | 1.125 | 0.055 | 421.2 | 5.6 | 242.4 | 0.218 | -12.1 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -12.2 | -1.5 | 25.48 | 21.57 | 1.181 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 378 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 162.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.24 | TEVA | 2022-04-26 10:33 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.55 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.08 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 52 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:29 | GW012-Mannul-SP30-MD-01 | 1.067 | 0.064 | 459.4 | 6.2 | 211.8 | 0.151 | -11.9 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.1 | MDx |  |  |  |  | -12 | -0.7 | 25.45 | 22.52 | 1.13 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 390 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 45.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.71 | TEVA | 2022-04-26 10:29 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.16 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.67 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-26 10:23 | GW012-Mannul-SP30-MD-01 | 1.044 | 0.066 | 459.4 | 6.2 | 223.7 | 0.163 | -12.3 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.7 | MDx |  |  |  |  | -12.4 | -1.6 | 25.49 | 22.68 | 1.124 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 344 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 27.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.46 | TEVA | 2022-04-26 10:23 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.43 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.36 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 0 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-25 15:12 | GW012-Mannul-SP30-MD-01 | [8 empty] | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx | [9 empty] | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 498 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 |  | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 |  | TEVA | 2022-04-25 15:14 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 |  | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 |  | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 60 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-25 15:08 | GW012-Mannul-SP30-MD-01 | 1.1 | 0.061 | 466 | 6.3 | 260.1 | 0.279 | -9.8 | -1.7 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.1 | MDx |  |  |  |  | -9.8 | -1.4 | 25.71 | 22.05 | 1.166 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 34.1 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.27 | TEVA | 2022-04-25 15:09 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.98 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.15 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-25 15:03 | GW012-Mannul-SP30-MD-01 | 1.082 | 0.071 | 431 | 5.8 | 297.6 | 0.188 | -9.7 | -2.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -9.7 | -2 | 25.18 | 21.75 | 1.158 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 498 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 46.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.05 | TEVA | 2022-04-25 15:03 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.23 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 74.64 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 68 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 2 | 120 | cfpharmtech\mi1chen | 2022-04-25 14:58 | GW012-Mannul-SP30-MD-01 | 1.119 | 0.042 | 456.3 | 6.2 | 252.2 | 0.086 | -11 | -2.2 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.3 | MDx |  |  |  |  | -11.1 | -2.1 | 25.71 | 22.48 | 1.144 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 394 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 176 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.18 | TEVA | 2022-04-25 14:59 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.51 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 75.92 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 102 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-001002#069 | SL1210805-05-05 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 14:43 | GW012-Mannul-SP30-MD-01 | 1.068 | 0.039 | 457 | 6.2 | 336.9 | 0.212 | -10.4 | -1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -10.2 | -1 | 25.36 | 22.7 | 1.117 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:43 | beclomethasone dipropionate | 001990-00 | 374 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 138.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.62 | TEVA | 2022-04-25 14:46 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.04 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.29 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-007#069 | SL1210805-05-05 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 14:39 | GW012-Mannul-SP30-MD-01 | 1.013 | 0.064 | 451.1 | 6.1 | 201.5 | 0.149 | -9.6 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx |  |  |  |  | -9.7 | -1.1 | 25.05 | 23.01 | 1.089 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:39 | beclomethasone dipropionate | 001990-00 | 388 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 46.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.86 | TEVA | 2022-04-25 14:46 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.56 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 77.1 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 84 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-044#069 | SL1210805-05-06 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 14:34 | GW012-Mannul-SP30-MD-01 | 1.204 | 0.069 | 485.3 | 6.6 | 258 | 0.111 | -12.1 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 37.9 | MDx |  |  |  |  | -12.1 | -1.5 | 27.72 | 22.29 | 1.243 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:33 | beclomethasone dipropionate | 001990-00 | 388 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 38 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 27 | TEVA | 2022-04-25 14:46 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.42 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 79.54 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 62 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-042#069 | SL1210805-05-02 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 14:30 | GW012-Mannul-SP30-MD-01 | [8 empty] | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38 | MDx | [9 empty] | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:29 | beclomethasone dipropionate | 001990-00 | 360 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 |  | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 |  | TEVA | 2022-04-25 14:46 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 |  | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 |  | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 56 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-008#069 | SL1210805-05-06 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 14:26 | GW012-Mannul-SP30-MD-01 | 1.133 | 0.051 | 476.6 | 6.5 | 208.6 | 0.136 | -12.9 | -1.3 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.3 | MDx |  |  |  |  | -13.1 | -1.2 | 27 | 22.5 | 1.2 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:25 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 178.5 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.93 | TEVA | 2022-04-25 14:26 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.89 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.24 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 92 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-066#069 | SL1210805-05-02 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 14:16 | GW012-Mannul-SP30-MD-01 | [8 empty] | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.3 | MDx | [9 empty] | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:15 | beclomethasone dipropionate | 001990-00 | 326 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 |  | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 |  | TEVA | 2022-04-25 14:21 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 |  | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 |  | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 70 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111206-043#069 | SL1210805-05-05 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 14:12 | GW012-Mannul-SP30-MD-01 | 1.101 | 0.055 | 412.4 | 5.6 | 323.7 | 0.074 | -11.2 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.6 | MDx |  |  |  |  | -11.2 | -1.5 | 24.67 | 21.26 | 1.16 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:11 | beclomethasone dipropionate | 001990-00 | 330 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 159.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.73 | TEVA | 2022-04-25 14:12 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.55 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 72.83 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 64 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-003#069 | SL1210805-05-05 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 14:09 | GW012-Mannul-SP30-MD-01 | 1.087 | 0.054 | 448.7 | 6.1 | 291.6 | 0.174 | -10.9 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.6 | MDx |  |  |  |  | -10.8 | -0.7 | 25.94 | 22.37 | 1.159 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:08 | beclomethasone dipropionate | 001990-00 | 396 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 155.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.62 | TEVA | 2022-04-25 14:09 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.65 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.28 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-068#069 | SL1210805-05-06 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 14:02 | GW012-Mannul-SP30-MD-01 | 1.073 | 0.06 | 477.9 | 6.5 | 234.8 | 0.212 | -11.8 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 38.7 | MDx |  |  |  |  | -11.9 | -0.7 | 26.27 | 22.48 | 1.169 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 14:02 | beclomethasone dipropionate | 001990-00 | 312 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 79.3 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.32 | TEVA | 2022-04-25 14:03 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.6 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.25 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 66 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 2 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-090138#069 | SL1210805-05-06 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 13:57 | GW012-Mannul-SP30-MD-01 | 1.113 | 0.06 | 461.7 | 6.3 | 272.7 | 0.062 | -10.8 | -1.1 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39 | MDx |  |  |  |  | -10.8 | -1 | 25.98 | 22.56 | 1.152 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:56 | beclomethasone dipropionate | 001990-00 | 406 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 162.7 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.3 | TEVA | 2022-04-25 13:57 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.72 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.98 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-002#069 | SL1210805-05-02 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 13:52 | GW012-Mannul-SP30-MD-01 | 1.07 | 0.043 | 480.8 | 6.5 | 247.1 | 0.094 | -9.9 | -0.9 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.3 | MDx |  |  |  |  | -10 | -0.8 | 26.12 | 23.04 | 1.134 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:51 | beclomethasone dipropionate | 001990-00 | 402 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 142.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.31 | TEVA | 2022-04-25 13:52 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.65 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 79.03 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 90 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111204-004#069 | SL1210805-05-06 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 13:47 | GW012-Mannul-SP30-MD-01 | 1.057 | 0.041 | 417.5 | 5.7 | 242.1 | 0.126 | -12.2 | -0.8 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 39.6 | MDx |  |  |  |  | -12.2 | -0.7 | 24.17 | 21.74 | 1.112 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:46 | beclomethasone dipropionate | 001990-00 | 384 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 149.8 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.4 | TEVA | 2022-04-25 13:47 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 22.13 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 73.15 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 88 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111203-002#069 | SL1210805-05-02 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 13:42 | GW012-Mannul-SP30-MD-01 | 1.098 | 0.06 | 472.7 | 6.4 | 191.6 | 0.183 | -12.7 | -1.5 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 40 | MDx |  |  |  |  | -12.9 | -1.5 | 26.46 | 23.01 | 1.15 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:42 | beclomethasone dipropionate | 001990-00 | 366 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 21.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 25.43 | TEVA | 2022-04-25 13:42 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.15 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 78.02 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 86 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | W21111205-067#069 | SL1210805-05-05 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 13:34 | GW012-Mannul-SP30-MD-01 | 1.028 | 0.055 | 455.2 | 6.2 | 226 | 0.339 | -12.5 | -1.6 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 40 | MDx |  |  |  |  | -12.8 | -1.3 | 25.18 | 22.84 | 1.102 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:34 | beclomethasone dipropionate | 001990-00 | 298 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 136.6 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 24.14 | TEVA | 2022-04-25 13:35 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 23.48 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 76.41 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 72 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 2 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen
GW012-RLD80-89 | AFS62A-038384#069 | SL1210805-05-02 | 1 | 120 | cfpharmtech\mi1chen | 2022-04-25 13:29 | GW012-Mannul-SP30-MD-01 | 1.06 | 0.044 | 405.1 | 5.5 | 243.4 | 0.1 | -12 | -1.4 | 001481-00 | 55056 | 001480-00 | 55066 | 5000 | 60 | 006379-00 |  | 40.3 | MDx |  |  |  |  | -12 | -1.3 | 23.94 | 21.23 | 1.128 | 2.8 GHz Xeon | 006218-00 | 1BFC2 | 14 | 14.5 | 006016-00 | 001910-00 | K15024414 | S1108696 |  | Manual | 2022-04-20 12:10 | Dell | Precision T5810 | 0.3 |  |  |  |  | 001142-00 | 1B26B0E | 2022-04-25 13:20 | beclomethasone dipropionate | 001990-00 | 412 |  | 0 | 500 | 250 | C | 400 | 001906-00 | 60000 | 006022-00 | 006019-00 | 6805CA3C67C5335CRE46981-008 | 51.2 | 0 | 2016-08-01 | 2 | NA | NA | 006684-00 | 7 | 154727 | 006017-00 | N/A | 001656-00 |  | RD215 | 23.08 | TEVA | 2022-04-25 13:30 | 2022-04-20 12:12 | cfpharmtech\mi1chen | SP30 1 SHOT | SprayVIEW SF MDI | 2 | 21.77 | 001714-00 |  | 001566-00 |  | Windows 7 Pro - SP1 | 30 | 71.86 | 006013-00 | EB57B36326 | 60-5000-0-400-0 | 8 GB | 5000 | 60 | cfpharmtech\mi1chen | 6H8RGB2 | [5 empty] | 001714-00 |  |  | 001566-00 |  | 8.1.1.4 | 94 | 5.3 | 2250 | 1BFEA | SprayVIEW SF MDI |  | 1 | Automatic | 000093-00 | 006219-00 |  | cfpharmtech\mi1chen

## Results-SP
Lot # | Identification # | Actuation Number | [5 empty] | Area [mm^2] | Area % | Ovality | Pattern Distance [mm] | Actuation Date | Method Name | Method Version
2 | 3 | 4 | [5 empty] | 11 | 12 | 35 | 100 | 7 | 8 | 93
Sample | Actuator | ActuationNumber_test | Product（Canister）Batch | Product（Canister) ID | Device Batch（MBA） | Device ID（MBA） | act | Area [mm^2] | Area_% | Ovality | Pattern Distance [mm] | Actuation Date | Method Name | Method Version
AFS62A-001002#069 | SL1210805-05-05 | 20 | AFS62A | 001002 | SL1210805 | 05-05 | 89 | 391.6 | 5.3 | 1.23 | 30 | 2022-05-04 13:11 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 20 | W21111205 | 068 | SL1210805 | 05-06 | 89 | 425.4 | 5.7 | 1.167 | 30 | 2022-05-04 13:08 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 20 | W21111204 | 004 | SL1210805 | 05-06 | 89 | 451.4 | 6.1 | 1.21 | 30 | 2022-05-04 13:06 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 20 | W21111206 | 042 | SL1210805 | 05-02 | 89 | 432.8 | 5.8 | 1.155 | 30 | 2022-05-04 13:03 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 20 | W21111204 | 002 | SL1210805 | 05-02 | 89 | 435.5 | 5.9 | 1.093 | 30 | 2022-05-04 13:00 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 20 | W21111203 | 008 | SL1210805 | 05-06 | 89 | 0 | 0 | 0 | 30 | 2022-05-04 12:56 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 20 | W21111206 | 043 | SL1210805 | 05-05 | 89 | 352.2 | 4.8 | 1.199 | 30 | 2022-05-04 12:53 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 20 | W21111205 | 067 | SL1210805 | 05-05 | 89 | 366.4 | 4.9 | 1.248 | 30 | 2022-05-04 12:50 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 20 | W21111203 | 002 | SL1210805 | 05-02 | 89 | 456.1 | 6.2 | 1.211 | 30 | 2022-05-04 12:47 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 20 | W21111205 | 066 | SL1210805 | 05-02 | 89 | 468.5 | 6.3 | 1.158 | 30 | 2022-05-04 12:44 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 20 | AFS62A | 090138 | SL1210805 | 05-06 | 89 | 448.7 | 6.1 | 1.144 | 30 | 2022-05-04 12:41 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 20 | W21111204 | 003 | SL1210805 | 05-05 | 89 | 370.1 | 5 | 1.164 | 30 | 2022-05-04 12:38 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 20 | W21111206 | 044 | SL1210805 | 05-06 | 89 | 451.3 | 6.1 | 1.207 | 30 | 2022-05-04 12:36 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 20 | W21111203 | 007 | SL1210805 | 05-05 | 89 | 367.5 | 5 | 1.134 | 30 | 2022-05-04 12:33 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 20 | AFS62A | 038384 | SL1210805 | 05-02 | 89 | 502.3 | 6.8 | 1.224 | 30 | 2022-05-04 12:30 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 19 | W21111206 | 044 | SL1210805 | 05-06 | 88 | 442 | 6 | 1.176 | 30 | 2022-05-04 12:28 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 19 | W21111204 | 002 | SL1210805 | 05-02 | 88 | 408.9 | 5.5 | 1.084 | 30 | 2022-05-04 12:25 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 19 | W21111205 | 066 | SL1210805 | 05-02 | 88 | 381.1 | 5.1 | 1.162 | 30 | 2022-05-04 12:22 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 19 | AFS62A | 090138 | SL1210805 | 05-06 | 88 | 427.8 | 5.8 | 1.193 | 30 | 2022-05-04 12:20 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 19 | W21111204 | 004 | SL1210805 | 05-06 | 88 | 448.5 | 6.1 | 1.159 | 30 | 2022-05-04 12:17 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 19 | W21111203 | 002 | SL1210805 | 05-02 | 88 | 437.7 | 5.9 | 1.147 | 30 | 2022-05-04 12:14 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 19 | W21111206 | 043 | SL1210805 | 05-05 | 88 | 305.7 | 4.1 | 1.121 | 30 | 2022-05-04 12:12 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 19 | W21111205 | 068 | SL1210805 | 05-06 | 88 | 471.2 | 6.4 | 1.179 | 30 | 2022-05-04 12:09 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 19 | W21111206 | 042 | SL1210805 | 05-02 | 88 | 384.5 | 5.2 | 1.191 | 30 | 2022-05-04 12:07 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 19 | W21111203 | 008 | SL1210805 | 05-06 | 88 | 411.6 | 5.6 | 1.149 | 30 | 2022-05-04 12:04 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 19 | W21111205 | 067 | SL1210805 | 05-05 | 88 | 304.1 | 4.1 | 1.208 | 30 | 2022-05-04 12:01 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 19 | AFS62A | 001002 | SL1210805 | 05-05 | 88 | 387.7 | 5.2 | 1.234 | 30 | 2022-05-04 11:59 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 19 | W21111203 | 007 | SL1210805 | 05-05 | 88 | 357 | 4.8 | 1.146 | 30 | 2022-05-04 11:56 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 19 | W21111204 | 003 | SL1210805 | 05-05 | 88 | 391.5 | 5.3 | 1.113 | 30 | 2022-05-04 11:54 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 19 | AFS62A | 038384 | SL1210805 | 05-02 | 88 | 465.7 | 6.3 | 1.198 | 30 | 2022-05-04 11:51 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 18 | W21111204 | 002 | SL1210805 | 05-02 | 87 | 372 | 5 | 1.131 | 30 | 2022-05-04 11:49 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 18 | W21111205 | 067 | SL1210805 | 05-05 | 87 | 372.7 | 5 | 1.083 | 30 | 2022-05-04 11:46 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 18 | W21111203 | 002 | SL1210805 | 05-02 | 87 | 411.5 | 5.6 | 1.14 | 30 | 2022-05-04 11:43 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 18 | W21111204 | 004 | SL1210805 | 05-06 | 87 | 433.7 | 5.9 | 1.099 | 30 | 2022-05-04 11:41 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 18 | W21111203 | 007 | SL1210805 | 05-05 | 87 | 376.1 | 5.1 | 1.219 | 30 | 2022-05-04 11:38 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 18 | W21111203 | 008 | SL1210805 | 05-06 | 87 | 412.1 | 5.6 | 1.122 | 30 | 2022-05-04 11:34 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 18 | W21111204 | 003 | SL1210805 | 05-05 | 87 | 289.8 | 3.9 | 1.153 | 30 | 2022-05-04 11:31 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 18 | W21111206 | 042 | SL1210805 | 05-02 | 87 | 434 | 5.9 | 1.147 | 30 | 2022-05-04 11:29 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 18 | W21111206 | 043 | SL1210805 | 05-05 | 87 | 346.5 | 4.7 | 1.089 | 30 | 2022-05-04 11:26 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 18 | AFS62A | 001002 | SL1210805 | 05-05 | 87 | 297.1 | 4 | 1.118 | 30 | 2022-05-04 11:23 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 18 | AFS62A | 090138 | SL1210805 | 05-06 | 87 | 407.8 | 5.5 | 1.141 | 30 | 2022-05-04 11:21 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 18 | W21111206 | 044 | SL1210805 | 05-06 | 87 | 431.1 | 5.8 | 1.104 | 30 | 2022-05-04 11:18 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 18 | W21111205 | 066 | SL1210805 | 05-02 | 87 | 389.2 | 5.3 | 1.09 | 30 | 2022-05-04 10:29 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 18 | W21111205 | 068 | SL1210805 | 05-06 | 87 | 463.4 | 6.3 | 1.136 | 30 | 2022-05-04 10:27 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 18 | AFS62A | 038384 | SL1210805 | 05-02 | 87 | 445.5 | 6 | 1.101 | 30 | 2022-05-04 10:24 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 17 | W21111205 | 067 | SL1210805 | 05-05 | 86 | 372.7 | 5 | 1.1 | 30 | 2022-05-04 10:21 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 17 | W21111204 | 004 | SL1210805 | 05-06 | 86 | 449.1 | 6.1 | 1.094 | 30 | 2022-05-04 10:19 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 17 | AFS62A | 090138 | SL1210805 | 05-06 | 86 | 387.1 | 5.2 | 1.162 | 30 | 2022-05-04 10:16 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 17 | W21111206 | 044 | SL1210805 | 05-06 | 86 | 392.4 | 5.3 | 1.208 | 30 | 2022-05-04 10:14 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 17 | W21111204 | 003 | SL1210805 | 05-05 | 86 | 424.9 | 5.7 | 1.153 | 30 | 2022-05-04 10:11 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 17 | W21111206 | 043 | SL1210805 | 05-05 | 86 | 392.1 | 5.3 | 1.234 | 30 | 2022-05-04 10:09 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 17 | W21111205 | 066 | SL1210805 | 05-02 | 86 | 431.3 | 5.8 | 1.258 | 30 | 2022-05-04 10:06 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 17 | AFS62A | 001002 | SL1210805 | 05-05 | 86 | 383.5 | 5.2 | 1.147 | 30 | 2022-05-04 10:03 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 17 | W21111206 | 042 | SL1210805 | 05-02 | 86 | 401.5 | 5.4 | 1.172 | 30 | 2022-05-04 10:01 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 17 | W21111203 | 007 | SL1210805 | 05-05 | 86 | 458.7 | 6.2 | 1.17 | 30 | 2022-05-04 09:58 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 17 | W21111205 | 068 | SL1210805 | 05-06 | 86 | 509.4 | 6.9 | 1.168 | 30 | 2022-05-04 09:55 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 17 | W21111204 | 002 | SL1210805 | 05-02 | 86 | 476.5 | 6.4 | 1.244 | 30 | 2022-05-04 09:52 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 17 | W21111203 | 008 | SL1210805 | 05-06 | 86 | 452.2 | 6.1 | 1.117 | 30 | 2022-05-04 09:49 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 17 | W21111203 | 002 | SL1210805 | 05-02 | 86 | 435 | 5.9 | 1.163 | 30 | 2022-05-04 09:47 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 17 | AFS62A | 038384 | SL1210805 | 05-02 | 86 | 426.2 | 5.8 | 1.204 | 30 | 2022-05-04 09:44 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 16 | W21111204 | 004 | SL1210805 | 05-06 | 85 | 385.1 | 5.2 | 1.147 | 30 | 2022-05-04 09:27 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 16 | AFS62A | 001002 | SL1210805 | 05-05 | 85 | 443.5 | 6 | 1.169 | 30 | 2022-05-04 09:24 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 16 | W21111206 | 044 | SL1210805 | 05-06 | 85 | 440.2 | 5.9 | 1.255 | 30 | 2022-05-04 09:22 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 16 | W21111203 | 007 | SL1210805 | 05-05 | 85 | 430.9 | 5.8 | 1.107 | 30 | 2022-05-04 09:19 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 16 | W21111203 | 008 | SL1210805 | 05-06 | 85 | 456.3 | 6.2 | 1.129 | 30 | 2022-05-04 09:16 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 16 | W21111206 | 042 | SL1210805 | 05-02 | 85 | 431.9 | 5.8 | 1.142 | 30 | 2022-05-04 09:14 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 16 | W21111204 | 003 | SL1210805 | 05-05 | 85 | 415.7 | 5.6 | 1.143 | 30 | 2022-05-04 09:11 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 16 | AFS62A | 090138 | SL1210805 | 05-06 | 85 | 519.1 | 7 | 1.116 | 30 | 2022-05-04 09:07 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 16 | W21111203 | 002 | SL1210805 | 05-02 | 85 | 508.1 | 6.9 | 1.105 | 30 | 2022-05-04 09:05 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 16 | W21111205 | 068 | SL1210805 | 05-06 | 85 | 454.5 | 6.1 | 1.169 | 30 | 2022-05-04 09:02 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 16 | W21111205 | 067 | SL1210805 | 05-05 | 85 | 537.5 | 7.3 | 1.17 | 30 | 2022-05-04 08:59 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 16 | W21111206 | 043 | SL1210805 | 05-05 | 85 | 460.7 | 6.2 | 1.125 | 30 | 2022-05-04 08:57 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 16 | W21111205 | 066 | SL1210805 | 05-02 | 85 | 434.6 | 5.9 | 1.155 | 30 | 2022-05-04 08:54 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 16 | W21111204 | 002 | SL1210805 | 05-02 | 85 | 476.8 | 6.4 | 1.185 | 30 | 2022-05-04 08:51 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 16 | AFS62A | 038384 | SL1210805 | 05-02 | 85 | 510.1 | 6.9 | 1.124 | 30 | 2022-05-04 08:48 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 15 | W21111203 | 007 | SL1210805 | 05-05 | 84 | 378.7 | 5.1 | 1.191 | 30 | 2022-04-29 16:28 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 15 | W21111205 | 066 | SL1210805 | 05-02 | 84 | 388.2 | 5.2 | 1.095 | 30 | 2022-04-29 16:26 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 15 | W21111205 | 067 | SL1210805 | 05-05 | 84 | 418.4 | 5.6 | 1.153 | 30 | 2022-04-29 16:23 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 15 | W21111204 | 003 | SL1210805 | 05-05 | 84 | 398.4 | 5.4 | 1.12 | 30 | 2022-04-29 16:20 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 15 | W21111204 | 002 | SL1210805 | 05-02 | 84 | 453.4 | 6.1 | 1.158 | 30 | 2022-04-29 16:17 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 15 | W21111203 | 008 | SL1210805 | 05-06 | 84 | 408.3 | 5.5 | 1.165 | 30 | 2022-04-29 16:15 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 15 | W21111204 | 004 | SL1210805 | 05-06 | 84 | 424.6 | 5.7 | 1.11 | 30 | 2022-04-29 16:12 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 15 | AFS62A | 001002 | SL1210805 | 05-05 | 84 | 486.1 | 6.5 | 1.174 | 30 | 2022-04-29 16:09 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 15 | W21111206 | 043 | SL1210805 | 05-05 | 84 | 434.5 | 5.9 | 1.114 | 30 | 2022-04-29 16:06 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 15 | AFS62A | 090138 | SL1210805 | 05-06 | 84 | 404.2 | 5.4 | 1.123 | 30 | 2022-04-29 16:04 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 15 | W21111205 | 068 | SL1210805 | 05-06 | 84 | 457.4 | 6.2 | 1.218 | 30 | 2022-04-29 15:59 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 15 | W21111203 | 002 | SL1210805 | 05-02 | 84 | 406.7 | 5.5 | 1.139 | 30 | 2022-04-29 15:56 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 15 | W21111206 | 042 | SL1210805 | 05-02 | 84 | 487.8 | 6.6 | 1.231 | 30 | 2022-04-29 15:54 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 15 | W21111206 | 044 | SL1210805 | 05-06 | 84 | 457.2 | 6.2 | 1.16 | 30 | 2022-04-29 15:50 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 15 | AFS62A | 038384 | SL1210805 | 05-02 | 84 | 408.1 | 5.5 | 1.119 | 30 | 2022-04-29 15:47 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 14 | W21111206 | 044 | SL1210805 | 05-06 | 83 | 425.6 | 5.7 | 1.178 | 30 | 2022-04-29 15:44 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 14 | W21111203 | 007 | SL1210805 | 05-05 | 83 | 428.8 | 5.8 | 1.195 | 30 | 2022-04-29 15:41 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 14 | W21111203 | 008 | SL1210805 | 05-06 | 83 | 432.9 | 5.8 | 1.15 | 30 | 2022-04-29 15:37 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 14 | W21111205 | 067 | SL1210805 | 05-05 | 83 | 446.5 | 6 | 1.192 | 30 | 2022-04-29 15:34 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 14 | W21111206 | 043 | SL1210805 | 05-05 | 83 | 384.8 | 5.2 | 1.173 | 30 | 2022-04-29 15:31 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 14 | W21111204 | 003 | SL1210805 | 05-05 | 83 | 389.6 | 5.2 | 1.163 | 30 | 2022-04-29 15:29 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 14 | W21111203 | 002 | SL1210805 | 05-02 | 83 | 379.3 | 5.1 | 1.122 | 30 | 2022-04-29 15:26 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 14 | AFS62A | 001002 | SL1210805 | 05-05 | 83 | 412.9 | 5.6 | 1.078 | 30 | 2022-04-29 15:23 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 14 | W21111205 | 066 | SL1210805 | 05-02 | 83 | 400 | 5.4 | 1.232 | 30 | 2022-04-29 15:21 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 14 | AFS62A | 090138 | SL1210805 | 05-06 | 83 | 401.2 | 5.4 | 1.13 | 30 | 2022-04-29 15:18 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 14 | W21111205 | 068 | SL1210805 | 05-06 | 83 | 442.4 | 6 | 1.143 | 30 | 2022-04-29 15:15 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 14 | W21111204 | 004 | SL1210805 | 05-06 | 83 | 494.1 | 6.7 | 1.065 | 30 | 2022-04-29 15:12 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 14 | W21111206 | 042 | SL1210805 | 05-02 | 83 | 435.6 | 5.9 | 1.117 | 30 | 2022-04-29 15:09 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 14 | W21111204 | 002 | SL1210805 | 05-02 | 83 | 461.3 | 6.2 | 1.255 | 30 | 2022-04-29 15:06 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 14 | AFS62A | 038384 | SL1210805 | 05-02 | 83 | 471.4 | 6.4 | 1.103 | 30 | 2022-04-29 15:04 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 13 | W21111204 | 003 | SL1210805 | 05-05 | 82 | 428.3 | 5.8 | 1.088 | 30 | 2022-04-29 15:00 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 13 | AFS62A | 090138 | SL1210805 | 05-06 | 82 | 489.5 | 6.6 | 1.258 | 30 | 2022-04-29 14:57 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 13 | W21111205 | 067 | SL1210805 | 05-05 | 82 | 459.8 | 6.2 | 1.133 | 30 | 2022-04-29 14:52 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 13 | AFS62A | 001002 | SL1210805 | 05-05 | 82 | 436.5 | 5.9 | 1.107 | 30 | 2022-04-29 14:47 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 13 | W21111204 | 004 | SL1210805 | 05-06 | 82 | 488.6 | 6.6 | 1.188 | 30 | 2022-04-29 14:43 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 13 | W21111206 | 043 | SL1210805 | 05-05 | 82 | 489.9 | 6.6 | 1.154 | 30 | 2022-04-29 14:40 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 13 | W21111203 | 007 | SL1210805 | 05-05 | 82 | 411.5 | 5.5 | 1.248 | 30 | 2022-04-29 14:36 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 13 | W21111206 | 042 | SL1210805 | 05-02 | 82 | 427.3 | 5.8 | 1.13 | 30 | 2022-04-29 14:33 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 13 | W21111203 | 008 | SL1210805 | 05-06 | 82 | 477.2 | 6.4 | 1.173 | 30 | 2022-04-29 14:30 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 13 | W21111205 | 068 | SL1210805 | 05-06 | 82 | 501.3 | 6.8 | 1.093 | 30 | 2022-04-29 14:28 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 13 | W21111206 | 044 | SL1210805 | 05-06 | 82 | 440.5 | 5.9 | 1.086 | 30 | 2022-04-29 14:25 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 13 | W21111205 | 066 | SL1210805 | 05-02 | 82 | 390.6 | 5.3 | 1.111 | 30 | 2022-04-29 14:22 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 13 | W21111203 | 002 | SL1210805 | 05-02 | 82 | 511.1 | 6.9 | 1.116 | 30 | 2022-04-29 14:20 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 13 | W21111204 | 002 | SL1210805 | 05-02 | 82 | 504 | 6.8 | 1.183 | 30 | 2022-04-29 14:16 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 13 | AFS62A | 038384 | SL1210805 | 05-02 | 82 | 445.8 | 6 | 1.075 | 30 | 2022-04-29 14:13 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 12 | W21111203 | 002 | SL1210805 | 05-02 | 81 | 449.3 | 6.1 | 1.095 | 30 | 2022-04-28 15:02 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 12 | W21111204 | 004 | SL1210805 | 05-06 | 81 | 480.4 | 6.5 | 1.235 | 30 | 2022-04-28 14:59 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 12 | W21111204 | 002 | SL1210805 | 05-02 | 81 | 405.8 | 5.5 | 1.244 | 30 | 2022-04-28 14:56 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 12 | AFS62A | 001002 | SL1210805 | 05-05 | 81 | 443.4 | 6 | 1.132 | 30 | 2022-04-28 14:53 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 12 | W21111203 | 007 | SL1210805 | 05-05 | 81 | 474.6 | 6.4 | 1.12 | 30 | 2022-04-28 14:50 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 12 | W21111204 | 003 | SL1210805 | 05-05 | 81 | 525.5 | 7.1 | 1.268 | 30 | 2022-04-28 14:48 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 12 | W21111205 | 068 | SL1210805 | 05-06 | 81 | 473.9 | 6.4 | 1.086 | 30 | 2022-04-28 14:45 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 12 | W21111205 | 066 | SL1210805 | 05-02 | 81 | 418.3 | 5.7 | 1.13 | 30 | 2022-04-28 14:43 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 12 | W21111206 | 043 | SL1210805 | 05-05 | 81 | 456.8 | 6.2 | 1.263 | 30 | 2022-04-28 14:40 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 12 | W21111205 | 067 | SL1210805 | 05-05 | 81 | 419.5 | 5.7 | 1.111 | 30 | 2022-04-28 14:37 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 12 | W21111206 | 042 | SL1210805 | 05-02 | 81 | 427.6 | 5.8 | 1.166 | 30 | 2022-04-28 14:33 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 12 | W21111203 | 008 | SL1210805 | 05-06 | 81 | 496.8 | 6.7 | 1.164 | 30 | 2022-04-28 14:30 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 12 | W21111206 | 044 | SL1210805 | 05-06 | 81 | 458 | 6.2 | 1.149 | 30 | 2022-04-28 14:27 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 12 | AFS62A | 090138 | SL1210805 | 05-06 | 81 | 461.2 | 6.3 | 1.125 | 30 | 2022-04-28 14:25 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 12 | AFS62A | 038384 | SL1210805 | 05-02 | 81 | 455.1 | 6.2 | 1.19 | 30 | 2022-04-28 14:22 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 11 | W21111203 | 008 | SL1210805 | 05-06 | 80 | 439.4 | 6 | 1.121 | 30 | 2022-04-28 14:19 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 11 | W21111203 | 007 | SL1210805 | 05-05 | 80 | 490.6 | 6.7 | 1.123 | 30 | 2022-04-28 14:17 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 11 | AFS62A | 090138 | SL1210805 | 05-06 | 80 | 450.6 | 6.1 | 1.104 | 30 | 2022-04-28 14:14 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 11 | W21111204 | 002 | SL1210805 | 05-02 | 80 | 499.2 | 6.8 | 1.104 | 30 | 2022-04-28 14:11 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 11 | W21111204 | 003 | SL1210805 | 05-05 | 80 | 460.1 | 6.2 | 1.19 | 30 | 2022-04-28 14:08 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 11 | W21111206 | 044 | SL1210805 | 05-06 | 80 | 435.9 | 5.9 | 1.152 | 30 | 2022-04-28 14:04 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 11 | W21111206 | 042 | SL1210805 | 05-02 | 80 | 490.1 | 6.7 | 1.215 | 30 | 2022-04-28 14:02 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 11 | W21111203 | 002 | SL1210805 | 05-02 | 80 | 477.6 | 6.5 | 1.118 | 30 | 2022-04-28 13:59 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 11 | W21111205 | 067 | SL1210805 | 05-05 | 80 | 491.5 | 6.7 | 1.115 | 30 | 2022-04-28 13:56 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 11 | W21111205 | 068 | SL1210805 | 05-06 | 80 | 454.6 | 6.2 | 1.188 | 30 | 2022-04-28 13:53 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 11 | W21111204 | 004 | SL1210805 | 05-06 | 80 | 487.4 | 6.6 | 1.201 | 30 | 2022-04-28 13:51 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 11 | W21111206 | 043 | SL1210805 | 05-05 | 80 | 436 | 5.9 | 1.108 | 30 | 2022-04-28 13:48 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 11 | W21111205 | 066 | SL1210805 | 05-02 | 80 | 443.9 | 6 | 1.187 | 30 | 2022-04-28 13:45 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 11 | AFS62A | 001002 | SL1210805 | 05-05 | 80 | 451 | 6.1 | 1.113 | 30 | 2022-04-28 13:42 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 11 | AFS62A | 038384 | SL1210805 | 05-02 | 80 | 428.4 | 5.8 | 1.135 | 30 | 2022-04-28 13:40 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 10 | W21111205 | 068 | SL1210805 | 05-06 | 79 | 471.1 | 6.4 | 1.169 | 30 | 2022-04-28 13:37 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 10 | W21111203 | 002 | SL1210805 | 05-02 | 79 | 407.1 | 5.5 | 1.116 | 30 | 2022-04-28 13:35 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 10 | W21111203 | 008 | SL1210805 | 05-06 | 79 | 431.9 | 5.9 | 1.167 | 30 | 2022-04-28 13:32 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 10 | W21111204 | 004 | SL1210805 | 05-06 | 79 | 483.9 | 6.6 | 1.198 | 30 | 2022-04-28 13:29 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 10 | W21111206 | 043 | SL1210805 | 05-05 | 79 | 454.6 | 6.2 | 1.135 | 30 | 2022-04-28 13:27 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 10 | W21111206 | 044 | SL1210805 | 05-06 | 79 | 426 | 5.8 | 1.132 | 30 | 2022-04-28 13:23 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 10 | W21111204 | 002 | SL1210805 | 05-02 | 79 | 417.7 | 5.7 | 1.113 | 30 | 2022-04-28 13:21 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 10 | W21111205 | 066 | SL1210805 | 05-02 | 79 | 426.2 | 5.8 | 1.128 | 30 | 2022-04-28 13:18 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 10 | AFS62A | 090138 | SL1210805 | 05-06 | 79 | 456.2 | 6.2 | 1.092 | 30 | 2022-04-28 13:15 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 10 | W21111205 | 067 | SL1210805 | 05-05 | 79 | 460.3 | 6.2 | 1.153 | 30 | 2022-04-28 13:13 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 10 | AFS62A | 001002 | SL1210805 | 05-05 | 79 | 381.9 | 5.2 | 1.205 | 30 | 2022-04-28 13:10 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 10 | W21111203 | 007 | SL1210805 | 05-05 | 79 | 430.1 | 5.8 | 1.112 | 30 | 2022-04-28 13:08 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 10 | W21111206 | 042 | SL1210805 | 05-02 | 79 | 460.1 | 6.2 | 1.17 | 30 | 2022-04-28 13:05 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 10 | W21111204 | 003 | SL1210805 | 05-05 | 79 | 449 | 6.1 | 1.14 | 30 | 2022-04-28 13:02 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 10 | AFS62A | 038384 | SL1210805 | 05-02 | 79 | 411.1 | 5.6 | 1.135 | 30 | 2022-04-28 12:59 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 9 | W21111204 | 004 | SL1210805 | 05-06 | 78 | 434.3 | 5.9 | 1.121 | 30 | 2022-04-28 11:20 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 9 | W21111205 | 068 | SL1210805 | 05-06 | 78 | 508.1 | 6.9 | 1.124 | 30 | 2022-04-28 11:18 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 9 | W21111204 | 002 | SL1210805 | 05-02 | 78 | 405.4 | 5.5 | 1.168 | 30 | 2022-04-28 11:15 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 9 | W21111203 | 008 | SL1210805 | 05-06 | 78 | 406.4 | 5.5 | 1.101 | 30 | 2022-04-28 11:13 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 9 | W21111205 | 066 | SL1210805 | 05-02 | 78 | 373.7 | 5.1 | 1.259 | 30 | 2022-04-28 11:10 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 9 | W21111203 | 002 | SL1210805 | 05-02 | 78 | 424.9 | 5.8 | 1.16 | 30 | 2022-04-28 11:07 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 9 | W21111203 | 007 | SL1210805 | 05-05 | 78 | 460.3 | 6.2 | 1.188 | 30 | 2022-04-28 11:05 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 9 | W21111204 | 003 | SL1210805 | 05-05 | 78 | 410.8 | 5.6 | 1.108 | 30 | 2022-04-28 11:02 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 9 | AFS62A | 090138 | SL1210805 | 05-06 | 78 | 415.1 | 5.6 | 1.131 | 30 | 2022-04-28 10:58 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 9 | W21111206 | 044 | SL1210805 | 05-06 | 78 | 391.3 | 5.3 | 1.133 | 30 | 2022-04-28 10:56 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 9 | W21111206 | 043 | SL1210805 | 05-05 | 78 | 431.6 | 5.9 | 1.114 | 30 | 2022-04-28 10:53 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 9 | W21111206 | 042 | SL1210805 | 05-02 | 78 | 418.5 | 5.7 | 1.138 | 30 | 2022-04-28 10:50 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 9 | W21111205 | 067 | SL1210805 | 05-05 | 78 | 488.1 | 6.6 | 1.174 | 30 | 2022-04-28 10:47 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 9 | AFS62A | 001002 | SL1210805 | 05-05 | 78 | 409.1 | 5.6 | 1.175 | 30 | 2022-04-28 10:44 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 9 | AFS62A | 038384 | SL1210805 | 05-02 | 78 | 442.8 | 6 | 1.205 | 30 | 2022-04-28 10:40 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 8 | W21111204 | 003 | SL1210805 | 05-05 | 77 | 484.9 | 6.6 | 1.194 | 30 | 2022-04-28 10:01 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 8 | W21111203 | 008 | SL1210805 | 05-06 | 77 | 390.9 | 5.3 | 1.12 | 30 | 2022-04-28 09:57 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 8 | W21111204 | 004 | SL1210805 | 05-06 | 77 | 447.7 | 6.1 | 1.203 | 30 | 2022-04-28 09:54 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 8 | W21111206 | 044 | SL1210805 | 05-06 | 77 | 436.2 | 5.9 | 1.223 | 30 | 2022-04-28 09:52 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 8 | W21111203 | 002 | SL1210805 | 05-02 | 77 | 450.5 | 6.1 | 1.141 | 30 | 2022-04-28 09:49 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 8 | W21111206 | 043 | SL1210805 | 05-05 | 77 | 389.3 | 5.3 | 1.134 | 30 | 2022-04-28 09:47 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 8 | W21111205 | 068 | SL1210805 | 05-06 | 77 | 435.3 | 5.9 | 1.18 | 30 | 2022-04-28 09:44 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 8 | W21111206 | 042 | SL1210805 | 05-02 | 77 | 417.8 | 5.7 | 1.186 | 30 | 2022-04-28 09:40 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 8 | W21111205 | 066 | SL1210805 | 05-02 | 77 | 405.4 | 5.5 | 1.228 | 30 | 2022-04-28 09:37 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 8 | W21111203 | 007 | SL1210805 | 05-05 | 77 | 476.7 | 6.5 | 1.137 | 30 | 2022-04-28 09:34 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 8 | AFS62A | 001002 | SL1210805 | 05-05 | 77 | 390 | 5.3 | 1.102 | 30 | 2022-04-28 09:31 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 8 | W21111204 | 002 | SL1210805 | 05-02 | 77 | 399.7 | 5.4 | 1.09 | 30 | 2022-04-28 09:28 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 8 | W21111205 | 067 | SL1210805 | 05-05 | 77 | 401.6 | 5.5 | 1.111 | 30 | 2022-04-28 09:25 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 8 | AFS62A | 090138 | SL1210805 | 05-06 | 77 | 379.6 | 5.2 | 1.107 | 30 | 2022-04-28 09:22 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 8 | AFS62A | 038384 | SL1210805 | 05-02 | 77 | 403.4 | 5.5 | 1.148 | 30 | 2022-04-28 09:19 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 7 | W21111203 | 008 | SL1210805 | 05-06 | 76 | 374.7 | 5 | 1.147 | 30 | 2022-04-27 16:57 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 7 | AFS62A | 090138 | SL1210805 | 05-06 | 76 | 403.7 | 5.4 | 1.1 | 30 | 2022-04-27 16:51 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 7 | W21111204 | 004 | SL1210805 | 05-06 | 76 | 0 | 0 | 0 | 30 | 2022-04-27 16:48 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 7 | AFS62A | 001002 | SL1210805 | 05-05 | 76 | 353.4 | 4.7 | 1.096 | 30 | 2022-04-27 16:45 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 7 | W21111204 | 003 | SL1210805 | 05-05 | 76 | 356.3 | 4.8 | 1.151 | 30 | 2022-04-27 16:42 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 7 | W21111205 | 066 | SL1210805 | 05-02 | 76 | 390 | 5.2 | 1.15 | 30 | 2022-04-27 16:39 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 7 | W21111206 | 042 | SL1210805 | 05-02 | 76 | 380 | 5.1 | 1.156 | 30 | 2022-04-27 16:34 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 7 | W21111205 | 067 | SL1210805 | 05-05 | 76 | 374.3 | 5 | 1.156 | 30 | 2022-04-27 16:31 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 7 | W21111206 | 043 | SL1210805 | 05-05 | 76 | 373 | 5 | 1.143 | 30 | 2022-04-27 16:16 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 7 | W21111204 | 002 | SL1210805 | 05-02 | 76 | 367.2 | 4.9 | 1.106 | 30 | 2022-04-27 16:13 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 7 | W21111206 | 044 | SL1210805 | 05-06 | 76 | 423.1 | 5.6 | 1.215 | 30 | 2022-04-27 16:10 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 7 | W21111203 | 007 | SL1210805 | 05-05 | 76 | 356 | 4.8 | 1.112 | 30 | 2022-04-27 16:07 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 7 | W21111205 | 068 | SL1210805 | 05-06 | 76 | 474.8 | 6.3 | 1.22 | 30 | 2022-04-27 16:03 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 7 | W21111203 | 002 | SL1210805 | 05-02 | 76 | 396.5 | 5.3 | 1.148 | 30 | 2022-04-27 16:00 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 7 | AFS62A | 038384 | SL1210805 | 05-02 | 76 | 317.7 | 4.2 | 1.125 | 30 | 2022-04-27 15:57 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 6 | W21111206 | 043 | SL1210805 | 05-05 | 75 | 430.6 | 5.7 | 1.109 | 30 | 2022-04-27 15:19 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 6 | AFS62A | 001002 | SL1210805 | 05-05 | 75 | 429.7 | 5.7 | 1.142 | 30 | 2022-04-27 15:16 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 6 | W21111203 | 002 | SL1210805 | 05-02 | 75 | 410.8 | 5.5 | 1.114 | 30 | 2022-04-27 15:14 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 6 | W21111203 | 008 | SL1210805 | 05-06 | 75 | 384.2 | 5.1 | 1.173 | 30 | 2022-04-27 15:11 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 6 | W21111205 | 067 | SL1210805 | 05-05 | 75 | 437 | 5.8 | 1.186 | 30 | 2022-04-27 15:07 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 6 | W21111204 | 003 | SL1210805 | 05-05 | 75 | 345.4 | 4.6 | 1.123 | 30 | 2022-04-27 15:04 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 6 | W21111204 | 002 | SL1210805 | 05-02 | 75 | 415.7 | 5.6 | 1.101 | 30 | 2022-04-27 15:00 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 6 | W21111206 | 044 | SL1210805 | 05-06 | 75 | 415.2 | 5.5 | 1.266 | 30 | 2022-04-27 14:57 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 6 | W21111205 | 066 | SL1210805 | 05-02 | 75 | 380.9 | 5.1 | 1.133 | 30 | 2022-04-27 14:54 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 6 | AFS62A | 090138 | SL1210805 | 05-06 | 75 | 388.5 | 5.2 | 1.218 | 30 | 2022-04-27 14:50 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 6 | W21111205 | 068 | SL1210805 | 05-06 | 75 | 478.9 | 6.4 | 1.145 | 30 | 2022-04-27 14:47 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 6 | W21111204 | 004 | SL1210805 | 05-06 | 75 | 402 | 5.4 | 1.112 | 30 | 2022-04-27 14:44 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 6 | W21111206 | 042 | SL1210805 | 05-02 | 75 | 399.4 | 5.3 | 1.2 | 30 | 2022-04-27 14:41 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 6 | W21111203 | 007 | SL1210805 | 05-05 | 75 | 432.1 | 5.8 | 1.093 | 30 | 2022-04-27 14:38 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 6 | AFS62A | 038384 | SL1210805 | 05-02 | 75 | 386.4 | 5.2 | 1.135 | 30 | 2022-04-27 14:35 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 5 | W21111205 | 066 | SL1210805 | 05-02 | 74 | 419.6 | 5.6 | 1.161 | 30 | 2022-04-27 14:25 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 5 | W21111203 | 002 | SL1210805 | 05-02 | 74 | 351.3 | 4.7 | 1.21 | 30 | 2022-04-27 14:22 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 5 | W21111203 | 008 | SL1210805 | 05-06 | 74 | 409.8 | 5.5 | 1.22 | 30 | 2022-04-27 14:19 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 5 | W21111206 | 044 | SL1210805 | 05-06 | 74 | 441.5 | 5.9 | 1.192 | 30 | 2022-04-27 14:16 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 5 | W21111206 | 042 | SL1210805 | 05-02 | 74 | 397.6 | 5.3 | 1.204 | 30 | 2022-04-27 14:13 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 5 | AFS62A | 001002 | SL1210805 | 05-05 | 74 | 386.4 | 5.2 | 1.123 | 30 | 2022-04-27 14:10 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 5 | W21111205 | 067 | SL1210805 | 05-05 | 74 | 461.2 | 6.2 | 1.102 | 30 | 2022-04-27 14:07 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 5 | W21111203 | 007 | SL1210805 | 05-05 | 74 | 413 | 5.5 | 1.137 | 30 | 2022-04-27 14:03 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 5 | W21111204 | 003 | SL1210805 | 05-05 | 74 | 404.8 | 5.4 | 1.092 | 30 | 2022-04-27 13:59 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 5 | W21111204 | 002 | SL1210805 | 05-02 | 74 | 432.4 | 5.8 | 1.104 | 30 | 2022-04-27 13:56 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 5 | AFS62A | 090138 | SL1210805 | 05-06 | 74 | 406.8 | 5.4 | 1.158 | 30 | 2022-04-27 13:53 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 5 | W21111205 | 068 | SL1210805 | 05-06 | 74 | 403.7 | 5.4 | 1.305 | 30 | 2022-04-27 13:50 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 5 | W21111206 | 043 | SL1210805 | 05-05 | 74 | 404.5 | 5.4 | 1.165 | 30 | 2022-04-27 13:47 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 5 | W21111204 | 004 | SL1210805 | 05-06 | 74 | 414.7 | 5.5 | 1.119 | 30 | 2022-04-27 13:44 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 5 | AFS62A | 038384 | SL1210805 | 05-02 | 74 | 404.4 | 5.4 | 1.123 | 30 | 2022-04-27 13:41 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 4 | W21111206 | 042 | SL1210805 | 05-02 | 73 | 447.6 | 6 | 1.237 | 30 | 2022-04-26 14:59 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 4 | W21111205 | 066 | SL1210805 | 05-02 | 73 | 378 | 5.1 | 1.198 | 30 | 2022-04-26 14:56 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 4 | AFS62A | 090138 | SL1210805 | 05-06 | 73 | 495.8 | 6.7 | 1.188 | 30 | 2022-04-26 14:53 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 4 | W21111203 | 007 | SL1210805 | 05-05 | 73 | 468.5 | 6.3 | 1.111 | 30 | 2022-04-26 14:50 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 4 | W21111205 | 067 | SL1210805 | 05-05 | 73 | 424.5 | 5.7 | 1.208 | 30 | 2022-04-26 14:47 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 4 | AFS62A | 001002 | SL1210805 | 05-05 | 73 | 452.9 | 6.1 | 1.193 | 30 | 2022-04-26 14:44 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 4 | W21111205 | 068 | SL1210805 | 05-06 | 73 | 473.8 | 6.4 | 1.14 | 30 | 2022-04-26 14:40 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 4 | W21111203 | 002 | SL1210805 | 05-02 | 73 | 472.8 | 6.4 | 1.123 | 30 | 2022-04-26 14:37 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 4 | W21111206 | 043 | SL1210805 | 05-05 | 73 | 495.5 | 6.7 | 1.175 | 30 | 2022-04-26 14:34 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 4 | W21111203 | 008 | SL1210805 | 05-06 | 73 | 452.6 | 6.1 | 1.112 | 30 | 2022-04-26 14:31 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 4 | W21111204 | 003 | SL1210805 | 05-05 | 73 | 430.5 | 5.8 | 1.123 | 30 | 2022-04-26 14:28 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 4 | W21111204 | 004 | SL1210805 | 05-06 | 73 | 433.1 | 5.8 | 1.153 | 30 | 2022-04-26 14:24 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 4 | W21111204 | 002 | SL1210805 | 05-02 | 73 | 412.7 | 5.6 | 1.109 | 30 | 2022-04-26 14:22 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 4 | W21111206 | 044 | SL1210805 | 05-06 | 73 | 451.1 | 6.1 | 1.217 | 30 | 2022-04-26 14:18 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 4 | AFS62A | 038384 | SL1210805 | 05-02 | 73 | 386.6 | 5.2 | 1.118 | 30 | 2022-04-26 14:15 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 3 | W21111204 | 004 | SL1210805 | 05-06 | 72 | 450.3 | 6.1 | 1.112 | 30 | 2022-04-26 14:12 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 3 | AFS62A | 090138 | SL1210805 | 05-06 | 72 | 421 | 5.7 | 1.11 | 30 | 2022-04-26 14:09 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 3 | W21111206 | 044 | SL1210805 | 05-06 | 72 | 476.7 | 6.4 | 1.187 | 30 | 2022-04-26 14:04 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 3 | W21111204 | 003 | SL1210805 | 05-05 | 72 | 453.8 | 6.1 | 1.245 | 30 | 2022-04-26 14:01 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 3 | W21111203 | 002 | SL1210805 | 05-02 | 72 | 410 | 5.5 | 1.147 | 30 | 2022-04-26 13:57 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 3 | W21111206 | 042 | SL1210805 | 05-02 | 72 | 435.6 | 5.9 | 1.118 | 30 | 2022-04-26 13:54 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 3 | W21111203 | 007 | SL1210805 | 05-05 | 72 | 502.6 | 6.8 | 1.263 | 30 | 2022-04-26 13:50 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 3 | AFS62A | 001002 | SL1210805 | 05-05 | 72 | 707.1 | 9.5 | 1.896 | 30 | 2022-04-26 13:43 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 3 | W21111206 | 043 | SL1210805 | 05-05 | 72 | 675.9 | 9.1 | 1.949 | 30 | 2022-04-26 13:39 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 3 | W21111204 | 002 | SL1210805 | 05-02 | 72 | 406.3 | 5.4 | 1.11 | 30 | 2022-04-26 13:23 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 3 | W21111205 | 067 | SL1210805 | 05-05 | 72 | 424.2 | 5.7 | 1.236 | 30 | 2022-04-26 13:19 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 3 | W21111205 | 066 | SL1210805 | 05-02 | 72 | 446.8 | 6 | 1.098 | 30 | 2022-04-26 13:14 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 3 | W21111203 | 008 | SL1210805 | 05-06 | 72 | 397.2 | 5.3 | 1.259 | 30 | 2022-04-26 13:09 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 3 | W21111205 | 068 | SL1210805 | 05-06 | 72 | 507.5 | 6.8 | 1.14 | 30 | 2022-04-26 13:06 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 3 | AFS62A | 038384 | SL1210805 | 05-02 | 72 | 387.5 | 5.2 | 1.155 | 30 | 2022-04-26 11:03 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 2 | W21111203 | 008 | SL1210805 | 05-06 | 71 | 440 | 5.9 | 1.135 | 30 | 2022-04-26 10:58 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 2 | W21111204 | 004 | SL1210805 | 05-06 | 71 | 409.3 | 5.5 | 1.117 | 30 | 2022-04-26 10:55 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 2 | W21111204 | 002 | SL1210805 | 05-02 | 71 | 463.2 | 6.2 | 1.129 | 30 | 2022-04-26 10:52 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 2 | AFS62A | 001002 | SL1210805 | 05-05 | 71 | 376.2 | 5 | 1.131 | 30 | 2022-04-26 10:49 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 2 | AFS62A | 090138 | SL1210805 | 05-06 | 71 | 443.5 | 5.9 | 1.173 | 30 | 2022-04-26 10:46 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 2 | W21111206 | 042 | SL1210805 | 05-02 | 71 | 413.7 | 5.5 | 1.174 | 30 | 2022-04-26 10:43 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 2 | W21111205 | 068 | SL1210805 | 05-06 | 71 | 450.9 | 6 | 1.114 | 30 | 2022-04-26 10:39 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 2 | W21111204 | 003 | SL1210805 | 05-05 | 71 | 435 | 5.8 | 1.157 | 30 | 2022-04-26 10:36 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 2 | W21111206 | 043 | SL1210805 | 05-05 | 71 | 421.2 | 5.6 | 1.181 | 30 | 2022-04-26 10:32 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 2 | W21111203 | 007 | SL1210805 | 05-05 | 71 | 459.4 | 6.2 | 1.13 | 30 | 2022-04-26 10:29 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 2 | W21111205 | 066 | SL1210805 | 05-02 | 71 | 459.4 | 6.2 | 1.124 | 30 | 2022-04-26 10:23 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 2 | W21111206 | 044 | SL1210805 | 05-06 | 71 | 0 | 0 | 0 | 30 | 2022-04-25 15:12 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 2 | W21111203 | 002 | SL1210805 | 05-02 | 71 | 466 | 6.3 | 1.166 | 30 | 2022-04-25 15:08 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 2 | W21111205 | 067 | SL1210805 | 05-05 | 71 | 431 | 5.8 | 1.158 | 30 | 2022-04-25 15:03 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 2 | AFS62A | 038384 | SL1210805 | 05-02 | 71 | 456.3 | 6.2 | 1.144 | 30 | 2022-04-25 14:58 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-001002#069 | SL1210805-05-05 | 1 | AFS62A | 001002 | SL1210805 | 05-05 | 70 | 457 | 6.2 | 1.117 | 30 | 2022-04-25 14:43 | GW012-Mannul-SP30-MD-01 | 2
W21111203-007#069 | SL1210805-05-05 | 1 | W21111203 | 007 | SL1210805 | 05-05 | 70 | 451.1 | 6.1 | 1.089 | 30 | 2022-04-25 14:39 | GW012-Mannul-SP30-MD-01 | 2
W21111206-044#069 | SL1210805-05-06 | 1 | W21111206 | 044 | SL1210805 | 05-06 | 70 | 485.3 | 6.6 | 1.243 | 30 | 2022-04-25 14:34 | GW012-Mannul-SP30-MD-01 | 2
W21111206-042#069 | SL1210805-05-02 | 1 | W21111206 | 042 | SL1210805 | 05-02 | 70 | 0 | 0 | 0 | 30 | 2022-04-25 14:30 | GW012-Mannul-SP30-MD-01 | 2
W21111203-008#069 | SL1210805-05-06 | 1 | W21111203 | 008 | SL1210805 | 05-06 | 70 | 476.6 | 6.5 | 1.2 | 30 | 2022-04-25 14:26 | GW012-Mannul-SP30-MD-01 | 2
W21111205-066#069 | SL1210805-05-02 | 1 | W21111205 | 066 | SL1210805 | 05-02 | 70 | 0 | 0 | 0 | 30 | 2022-04-25 14:16 | GW012-Mannul-SP30-MD-01 | 2
W21111206-043#069 | SL1210805-05-05 | 1 | W21111206 | 043 | SL1210805 | 05-05 | 70 | 412.4 | 5.6 | 1.16 | 30 | 2022-04-25 14:12 | GW012-Mannul-SP30-MD-01 | 2
W21111204-003#069 | SL1210805-05-05 | 1 | W21111204 | 003 | SL1210805 | 05-05 | 70 | 448.7 | 6.1 | 1.159 | 30 | 2022-04-25 14:09 | GW012-Mannul-SP30-MD-01 | 2
W21111205-068#069 | SL1210805-05-06 | 1 | W21111205 | 068 | SL1210805 | 05-06 | 70 | 477.9 | 6.5 | 1.169 | 30 | 2022-04-25 14:02 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-090138#069 | SL1210805-05-06 | 1 | AFS62A | 090138 | SL1210805 | 05-06 | 70 | 461.7 | 6.3 | 1.152 | 30 | 2022-04-25 13:57 | GW012-Mannul-SP30-MD-01 | 2
W21111204-002#069 | SL1210805-05-02 | 1 | W21111204 | 002 | SL1210805 | 05-02 | 70 | 480.8 | 6.5 | 1.134 | 30 | 2022-04-25 13:52 | GW012-Mannul-SP30-MD-01 | 2
W21111204-004#069 | SL1210805-05-06 | 1 | W21111204 | 004 | SL1210805 | 05-06 | 70 | 417.5 | 5.7 | 1.112 | 30 | 2022-04-25 13:47 | GW012-Mannul-SP30-MD-01 | 2
W21111203-002#069 | SL1210805-05-02 | 1 | W21111203 | 002 | SL1210805 | 05-02 | 70 | 472.7 | 6.4 | 1.15 | 30 | 2022-04-25 13:42 | GW012-Mannul-SP30-MD-01 | 2
W21111205-067#069 | SL1210805-05-05 | 1 | W21111205 | 067 | SL1210805 | 05-05 | 70 | 455.2 | 6.2 | 1.102 | 30 | 2022-04-25 13:34 | GW012-Mannul-SP30-MD-01 | 2
AFS62A-038384#069 | SL1210805-05-02 | 1 | AFS62A | 038384 | SL1210805 | 05-02 | 70 | 405.1 | 5.5 | 1.128 | 30 | 2022-04-25 13:29 | GW012-Mannul-SP30-MD-01 | 2